## DR X-ray
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | General Purpose X-ray Machine (DR X-ray) | 1 | Comply Yes/No (If Yes Specify) | Supporting Documents Ref. | by Hospital
Specifications
The start -of -the-Art Digital Radiography (DR X-ray) system with maximum 80-90 kW ,900 mA,FDP 17x17 2-numbers, Long length detector with 99-140 micron,To minimize X-ray dose and improves Examination workflow , Single Auto Stitching ,Quick preview after exposure ,Ceiling/wall mounted ,Tilting Bucky and Auto positioning features.
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | All pre-installation requirements should be submitted along with the quotation including suggested room dimensions etc.
4 | Integration with PACS , GE Dose Watch & RIS :Unit must be able to integrate with hospital PACS system & GE Dose Watch .All software's and hardware's must be provided by vendor to do the integration .
4.1 | DICOM Image Storage Service Class as defined below: The DICOM Storage Commitment Service Class For retrieving prior images from the PACS, the DICOM Query/Retrieve (C-FIND and C-MOVE for Patient Root and/or Study Root Model) Service Class should be support.
4.2 | Fully digital radiography system with Indirect conversion TFT technology, capable of detector exposure in vertical, horizontal and oblique positions to perform general radiography. The unit should be completely integrated (integrated Generator and Image Acquisition) and comprise the following along with auto quality control features incorporated.
5 | X-ray Tube & Suspension:
5.1 | Dual focus X-ray tube with high anode heat capacity and cooling rate
5.2 | Output power of 80-90 KW
5.3 | Telescopic ceiling suspension with motorized height adjustment
5.4 | Control handle with automatic collimation with timer
5.5 | Collimation light switches automatically on when needed
5.6 | SID positioning with laser beam & tape line
5.7 | Digital SID display
5.8 | Protection Device like Broken filament, Abnormal rotor, Overload and High voltage arcing
5.9 | The mAs are displayed according to the constant chosen an optional display of the mAS accumulates during the exposure the mAs are also displayed during an exposure when using the automatic exposure controller
5.1 | Error code display: 1.Tube overheating. 2.High voltage, power supply not ready. 3.Exposure request which exceeds the max power the generator exposure request for too many mAs for the exposure. 4.The exposure is impossible due to the absence of the emission characteristics of the tube
6 | Control & Image Pre-Processing Console
6.1 | Ceiling mounted type
6.2 | Compact control console with display screen
6.3 | Organ programming features
6.4 | Manual & automatic parameter settings
6.5 | Automatic exposure control (AEC) 3 chambers
6.6 | Post exposure data & Under exposure indication
6.7 | Tube protection against over load
6.8 | Anatomically specific pre-processing
6.9 | Auto ranging feature
6.1 | Integrated DAP Monitor with shortest warm-up time
7 | Exposure Index feature
7.1 | Archival Storage: Local Hard Drive capacity > 1 Tera Byte, RW/+RW/-DL/Read, CD-R, DVD,+RW support for Image storage, ≥23’’ HR flat panel 5MP High resolution medical grade monitor with status and error messages on in-built/1 external PC monitor
7.2 | Fully customizable patient work list
7.3 | Software as required with no additional cost
7.4 | Upgradeable to higher version
7.5 | The no. of mouse clicks to be performed to get an image should be as minimum as possible, but should not exceed 3 mouse clicks
7.6 | Touch screen for fast and intuitive operation in a dynamic X- ray environment
7.7 | Full screen viewing mode for fast decisions
7.8 | 1, 2 , 3 to 4 image display for image comparison
7.9 | The patient name should disappear from the patient display automatically after the first exposure for privacy reasons
7.1 | Auto/Manual Pre-filter selection
7.11 | A separate workstation for image positioning and patient demographic data is required and the workstation should be able to send and receive DICOM (Digital Imaging and Communications in Medicine) standards
8 | Automatic Tracking, sensing and alignment
8.1 | Automatic Positioning for X-ray Tube & Detectors
8.2 | Reset to last SID position
8.3 | Electromagnetic brakes in all directions & central release feature
8.4 | Tube Heat dissipation to be specified
8.5 | Focal spot to be specified
8.6 | Anode heat storage to be specified
8.7 | Unlimited focal spot sizes using a Mixed focal spot of large & small focal for sharper images at high generator power, allowing Better utilization of both cathodes, improving lifetime of cathode
8.8 | Facility to release all movements by click of a single button
8.9 | Automatic SID measurement tool
8.1 | Pre-set value indication for KV, mA, S and mAs for viewing and reading
9 | Table Hight adjustable:
9.1 | Height adjustment range min 40 cm to 90 cm
9.2 | Specify Table top dimensions
9.3 | Specify Table top travel range : 1)longitudinal 2)Transverse
9.4 | Max. patient load ≥300 Kgs
9.5 | Balanced at counter weight and table movement: 4 ways with breaks
9a | Digital Bucky Table:
9.1 | Motorized height adjustment
9.2 | Floating table top
9.3 | Load bearing capacity up to 250Kg or more
9.4 | Collision prevention sensors for downward movement
9.5 | CSI Digital Flat Panel Detector with highest DQE - Fixed Type Detector
9.6 | Minimum Detector size 17X17 inches
9.7 | Spatial Resolution 3.5 LP/mm or higher ( must specify in the datasheet) Oscillating, easily removable grid
9.8 | Cassette sensing feature
9.9 | Easy to clean surface
9.1 | Foot & hand switch for operation
9.11 | Emergency power off switch
9.12 | Table Smooth movement: 4 ways with breaks
10 | Digital Wall Bucky Stand:
10.1 | Height adjustable and Tiltable (-20 to +90)
10.2 | Specify vertical movement travel range
10.3 | CSI Digital Flat Panel Detector with highest DQE and antiscatter grid - Fixed Type Detector
10.4 | Pixel size: <140 µm
10.5 | Large Active Image Size, 43 x 43 cm or more Spatial Resolution 3.5 LP/mm or higher (must specify in the datasheet)
10.6 | Oscillating easily removable grid
10.7 | Cassette sensing feature(Automatic tube tracking )
10.8 | wired or Wireless remote control
10.9 | Easy to clean surface
10.1 | with automatic exposure control
11 | Software:
11.1 | Repeat /Reject Analysis software/DAP
11.2 | Statistical tool software
11.3 | Protocol analysis
12 | Unit must be antimicrobial Protection
13 | Specify Weight and Dimensions of the unit
14 | Power consumption & Heat dissipation in BTU/h
15 | Accessory:
15.1 | All start-up accessories as per end user to be included
15.2 | Half Mobile Lead apron on wheels x 1
15.3 | Half lead apron x 1
15.4 | Radiology Positioning Devices -Full Set
15.5 | Child Immobilizers and Pediatric Positioning Complete Unit
15.6 | Weight Bearing Steps(Single and Double stands) & Hand Holders -2 sets
15.7 | Positioning Sponges / Sandbags with holders -Each 2 no's
15.8 | Standard Radiation Protection ;Lead Apron 2 nos,Thyroid Shield 2 no's
15.9 | Immobilization Straps -full set 2 no's each
15.1 | Whole Spine Imaging accessories -Each 2 no's
16 | Trade-in (If relevant)
16.1 | Supplier should offer a Trade-in price for the following existing system:
16.2 | DESCRIPTION MANUFACTURER MODEL YEAR OF PRUCHASE
16.3 | GENERAL RD SYSTEM SIEMENS, GERMANY AAXIOM ARISTOS MX/VX 21/07/2011
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
17.4 | Quality control ,Quality assurance and acceptance test should be done by the company
17.5 | Safety assessment should be done by the company
18 | To provide information on energy efficiency options in the equipment, sustainability methods adopted, any carbon footprint technologies utilized, and certifications
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Available Additional Option
| OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
21 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Radiology Technicians and doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 10 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at least 3 years comprehensive from date of installation and equipment acceptance.
12 | Vendor Must include Site preparation (pre-Installation) work for the proposed model

## Pre-Installation Works
Civil and MEP works for General Purpose X-ray Machine
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Civil and MEP works for Fluoroscopy unit | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
All Pre-installation requirements should be submitted, including suggested room Design and dimensions
1 | All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required.
2 | Scope of Proposed Design and As built Drawings ( after completion of Project)Should be Included
3 | Interior Designing and Decorative works to be Proposed and should Include in scope of Works
4 | Demolition and Mobilization works should be included if required
5 | Mobilization(Wall opening and closing ) for X-ray machine delivery should be included
6 | The system offered shall be designed to operate normally under the conditions of UAE, Environmental factors, FNAR and MOH policy the conditions include Power Supply,Climate,Temperature, Humidity,Dust,etc
7 | Detailed Lead test reports for existing lead and compliance certification to be provided at each step to ensure the radiation leakage not only meets but also exceeds the manufacturers requirements.
8 | Provide overlap of the material at each panel joint for the maximum level of performance, attenuation integrity and electrical conductivity.
9 | To install appropriate internal fit out or lining, lighting and electrical meet international standards.
10 | supply and installation of Warning light and signs
11 | Partition works and Ceiling works should be Similar to UHS standards. Proposed Specification need to be provided
12 | Aesthetic Ceiling picture can be done in X-ray room as per client suggestion
13 | Acoustic ceiling tiles.
14 | The Sound insulation can be incorporated into the framework of the walls during assembly and to ensure standard acoustic internal lining.
15 | Wall treatment
16 | To provide clearest observation window with observation glass
17 | The observation Glass shall be anti-reflective coating
18 | Flooring- Medical Grade (similar or better than existing flooring at UHS)
19 | Removing of existing flooring & levelling with self-levelling screed for complete area.
20 | Power Supply Capacity and Specifications
21 | Proposal for Supply and Installation of Electrical works. socket and switches should be via brand and of same type used in UHS premises.(All wires will be concealed and in conduct pipes and the wiring for electrical and network cable will be commensurate with the existing ).
22 | The lighting should be all LED and should support energy conservation
23 | HVAC- Study the current layout and Propose the modifications required
24 | Fire Alarm and Fire Fighting modification works to be propose
25 | Supply and installation Bedhead units for proposed locations including Medical Gas Points and Nurse call system
26 | Proposal for Supply and Installation of 4mm laed lining up to 3 mters from floor lavel as per FANR norms.
27 | Existing ceiling mounted procedure light and ceiling mounted lead shield can be used.
28 | All MEP work including medical gas system should be according to UHS stander which included future line if needed and cureent system assessments to avoid any strugeling with operation after the installation
29 | lead lining should be according to the FANR requirements
Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
30 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Equipment. Two Copies each and one Soft Copy for each.
3 | It is mandatory for the Contractor to submit End-of-Sale, End-Of-Support, and End-Of-Life for each individual hardware component - Proof documents from the manufacturer to be attached with the proposal.
4 | Proposal Should be submitted after having site survey and comparison with the current As built
5 | All The works Should Comply with Standards of Civil defense, SEWA and SHA
6 | Contractor should discuss the final technical proposal with the technical team before submission. Proposal should include ongoing warranty, support.