## COMPUTED TOMOGRAPHY (CT)SYSTEM
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | COMPUTED TOMOGRAPHY (CT)SYSTEM | 1 | Comply Yes/No (If Yes Specify) | Supporting Documents Ref. | by Hospital
Specifications
The state -of -the-Art model designed for high patient throughput and superior image quality. Actual 256 Slice CT Whole Body Scanner with advanced image reconstruction capabilities,360 degree rotation. 90-100 kW Generator ,900 mA or higher ,4-8 cm detectors size, Integrated AI-Based Solutions for automated image analysis, workflow optimization. Advanced Low-Dose Technology ensuring optimal diagnostic performance with minimal radiation exposure, including General & Trauma clinical applications.
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | All pre-installation requirements should be submitted along with the quotation including suggested room dimensions etc.
4 | Integration with PACS , GE Dose Watch & RIS :Unit must be able to integrate with hospital PACS system & GE Dose Watch .All software's and hardware's must be provided by vendor to do the integration .
4.1 | DICOM Image Storage Service Class as defined below: The DICOM Storage Commitment Service Class For retrieving prior images from the PACS, the DICOM Query/Retrieve.
5 | Specify Make/Model/Origin/Year of release
6 | Gantry:
6.1 | Should have user friendly control panels for easy positioning.
6.2 | Automation patient position and reconstruction
6.3 | Gantry should be incorporated with Low voltage slip rings
6.4 | Slip ring: Optical or RF with higher data transfer rate
6.5 | 360° Rotation time : less than .35 sec (lowest time is preferred)
6.6 | Best temporal resolution
6.7 | The gantry should have a minimum tilt of 30 degrees on either side
6.8 | Quiet operation while acquisition at higher rotational speed
6.9 | Aperture should be at least 70 cm diameter or higher
6.1 | Smart Solution, integrated camera for patient positioning and workflow to be offered for patient
6.11 | Scan Localizer: Laser beam
6.12 | Bilateral touch panel control for positioning lights & table control
6.13 | Two way intercom between gantry & console
6.14 | Integrated color HR video camera to monitor patient for auto position (To be quoted separately as a option)
6.15 | Required ECG HW & SW
6.16 | Fluoro & Angio Hardware
6.17 | Inside monitor for fluoroscopy
6.18 | Large Display for Fluoro/Angio with properly mounted, or integrated ( to be quoted separately)
6.19 | All the gantry cooling devices should be inbuilt only
6.2 | Details of SFOV (Scanner Field of View), Reconstructed Field of View and the Display Field of View (DFOV)
6.21 | Pitch is to be free AUTO mode and option for manual mode and to specify the ranges
6.22 | Emergency stop on gantry also on scan room wall
7 | Patient Table:
7.1 | Floating Carbon fiber table top
7.2 | Should have a minimum load bearing capacity of at least 250 kg
7.3 | Should have Bariatric patient support up to 300 Kg
7.4 | Horizontal Scan Range: 170 cm or more
7.5 | Couch controls on console
7.6 | Foot switch
7.7 | Fitted with collision detectors & patient straps
7.8 | IV pole
8 | Detector:
8.1 | Detector type: Solid state array with 64 rows , 128 rows (physical detector rows)(Higher is preferred)
8.2 | Maximum coverage in one rotation
8.3 | Standard Field of view: 50 cm or more
8.4 | Number of simultaneous acquired slices/rotation: 256 slices with highest detector width range ( Larger is preferred) Please specify the number of detector elements and row including Z axis,2) Detector should have facility to acquire 256 slices simultaneously in one rotation
8.5 | Largest Detector coverage area of minimum 8 cm and Please specify Z-axis whole width
8.6 | The effective elements /channels should be at least 650 per row
8.7 | Spatial Image resolution: app. from 12 lp/cm and higher and the minimum thickness should be 0.625 mm or less with better image,
8.8 | Digital real time filtering for advanced noise reduction with motion detection
8.9 | The Number of views/rotation at fastest scan mode details
8.9 | Integrated DAP meter
9 | Generator & X ray- tube:
9.1 | High frequency, 3 phase, microprocessor controlled
9.2 | Power output ≥ 100 KW
9.3 | kV range : variable up to 140 kV or more
9.4 | Tube focal spot: Dual focus
9.5 | Tube anode heat storage capacity > 6 MHU and More
9.6 | Tube cooling rate > 700 kHU/min
9.7 | Tube cooling: Air/Water or Air/Oil cooled (all pre installation requirement to be done by supplier, in case if required)
9.8 | mA range: app 10-800 mA or more in increments of maximum 10 mA
9.9 | Optimum patient dose management
9.1 | Lower KV options for paediatric patients
9.11 | Dose reduction techniques/features
10 | Image Processing:
10.1 | High-performance computer
10.2 | Reconstruction matrix: minimum 512x512
10.3 | Reconstruction rate: ≥ 50 image/sec at 512 x 512 image matrix (Display Matrix)
10.4 | Scan field of view up to 50 cm or more continuous
10.5 | Best metal Artifact Reduction System
10.6 | Real time reconstruction speed should be maximum possible please specify
10.7 | Reconstructed slice thickness should be up to 5mm
10.8 | Image processing, viewing and storage software
11 | Console Display:
11.1 | High resolution Monitor size: Two no’s fusion Medical Grade(6 MP) minimum 27” HR flat panel or higher and RAM of 16 GB and the latest multitasking computer should be offered with latest processors.
11.2 | Image Display Matrix: 1920 x 1200 or more
11.3 | Multi-image display and Computer desk and cabinet to be included
12 | Image Storage:
12.1 | The Hard Disk capacity for both image and raw data should be 2 TB or more and system to have facility to store at least 500,000 images
12.2 | Archival Storage: RW/+RW/-DL/Read, CD-R, DVD,+RW support, USB
13 | Software to be specified if standard or available and optional (at the console it self not on the satellite workstation)
13.1 | Multi-planar Reconstruction (MPR) of secondary views, with viewing perspectives in all planes including curved and orthogonal MPR
13.2 | Complete software for reconstruction,3D Surface Shaded Display,rendring and volume measurement
13.3 | Metal Artifact reduction software and Details of stitching Artifacts
13.4 | Dose reduction techniques/features
13.5 | Cardiac Imaging
13.6 | Neuro Imaging
13.7 | Spectral Imaging
13.8 | Comprehensive cardiac package (including Calcium Scoring and prospective & retrospective gating)
13.9 | Advanced Cardiovascular Imaging
13.1 | Dynamic Myocardial stress perfusion imaging
13.11 | Real time Fluoro with all required hardware & software
13.12 | Automatic Multiphase Angio CT
13.13 | Vascular Interventional
13.14 | Biopsy software
13.15 | Perfusion CT imaging
13.16 | Functional CT imaging
13.17 | Stroke Imaging
13.18 | Advanced Neuro package
13.19 | 3D Volume Rendering, Max Intensity projection (MIP)
13.2 | Auto bone removal
13.21 | Virtual endoscopy package, Virtual Colonoscopy, Virtual bronchoscopy
13.22 | Pulmonary Package
13.23 | Pediatric imaging as standard
13.24 | Liver volume measurement
13.25 | Scanner –Injector coupling including injector protocol management
13.26 | Dedicated high-end workstation with High Resolution Diagnostic large Dual Displays with Full DICOM 3 and integrated with PACS existing system
14 | Dose Management system:
14.1 | Software and Hardware for complete Dose management system for adult and pediatric(Mandatory)
15 | Networking:
15.1 | DICOM 3 conformant with licenses for DICOM Storage SCU/SCP, Query/Retrieve SCU, Modality work list SCU, Storage Commitment SCU, Basic grayscale Print SCU, CD-R Data exchange, & Verification SCU/SCP, MPPS etc.
15.2 | ECG Waveform SCP/SCU
15.3 | Full Integration with existing PACS
16 | UPS:
16.1 | UPS capable of supporting the FULL system for up to 30 min must be included
16.2 | Offered CT should be compatible fully integrating with existing Contrast Medical Injector , all cost of integration are to be included in the offer, or offer equivalent Injector compatible with excising Consumables Contract, all preinstallation of Injector are to be included subject for approval by End-users (Dual head pressure injector with auto trigger of scan at peak opacification should be provided)
17 | Software:
17.1 | Repeat /Reject Analysis software
17.2 | Statistical tool software
17.3 | Protocol analysis
17.4 | Unit must be antimicrobial Protection-Easy to clean surface
17.5 | Specify Weight and Dimensions of the unit
17.6 | Power consumption & Heat dissipation in BTU/h
18 | Accessory:
18.1 | All start-up accessories as per end user to be included
18.2 | Half Mobile Lead apron on wheels x 1 ,Patient foot rest, shoulder supports, patient hand grips, compression band, Armpit support.
18.3 | Half lead apron x 1
18.4 | Whole Spine Imaging accessories
18.5 | Mobile Detector holder ,bed holder
19 | Trade-in (If relevant)
19.1 | Supplier should offer a Trade-in price for the following existing system:
19.2 | DESCRIPTION MANUFACTURER MODEL YEAR OF PRUCHASE
19.3 | COMPUTED TOMOGRAPHY SYSTEM SIEMENS, GERMANY SOMATOM DEFINITION AS 2011
20 | Quality and Patient safety
20.1 | Unit must be FDA / CE certified
20.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
20.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
20.4 | Safety assessment should be done by the company to know the lead thickness in mm
20.5 | Quality control ,Quality assurance and acceptance test should be done by the company after installation of Equipment
20.6 | supplier must provide the information on energy efficiency options in the equipment, sustainability methods adopted, any carbon footprint technologies utilized, and certifications
21 | REQUIREMENTS FOR SERVER BASED MULTIMODALITY WORKSTATION SOLUTION FOR IMAGING APPLICATIONS - Work-Station for advanced evaluation and reporting
21.1 | 1)The work-station should be completely and seamlessly integrated with existing DICOM PACS system in the department with below mentioned applications made available on the existing users 2) The work-station should have the latest quadcore / octacore processor with at least 32 GB RAM and 2TB storage with concurrent processing capability of at least 20000 images and Dedicated high-end medical grade two monitor 6 mp not less than 27" inch. 3)Full soft ware Applications ( : a. Lung nodule with CAD and follow up with tumour tracking calculation. - 1 concurrent license b. Lung analysis with emphysema evaluation. - 1 concurrent license c. Comprehensive liver analysis with auto liver segmentation, vascular region based lobe segmentation, lesion auto contouring, tumour burden, calculation.,Cardiac Software with spcify the Type and the Level of iterative reconstruction software -- 1 concurrent license d. CT body, whole organ and whole brain long range perfusion. 1 concurrent license e. Multimodality fusion.- 1 concurrent license
22 | OTHER TERMS:
22.1 | 1. Supplier should assist with necessary diagrams and documents for Radiation safety department approval. 2. Set of lead aprons to be provided . 3. All patient accessories should including head rest and patient restrainers should be included as 2 sets. 4. Ceiling mounted diagnostic quality slave monitor with functional integration to console for performing CT guided interventions — 1 5. Patient support . 6. 1000 sets of syringes and tubing, to be used along with the pressure injector. 7. Temperature controlled smart cabinet (Medium size)for storage of contrast vials — 1 8. Multiparametric monitor for measurement of vitals and SpO2 — 1 9. Procedure trolley -1 10. Height adjustable revolving chairs with rubberized swivel wheels and armrest (of reputed make / vendor) - 5 numbers 11. Multilevel shifter trolley with facility for oxygen cylinder — 1 12. Collapsible, wheel chair with rubberized swivel wheels of standard & reputed make — 1 13. Digital BP apparatus of reputed make -1 14. Contrast Warmer -1 15. Patient positioning cushions -2 set 16. Ceiling lights with dimmer. 17. Operators & technical manuals of main and subsystems, Cooling chart curve of the x- ray tube and all other relevant material, if any, like concept guide/ books; notable published work must be supplied in duplicate 18. Training to be done for 10days in Three to four phase for a different batch. 19) Advanced training — 1 Radiologists, 2 Technologists and 1 biomedical Engineer (nominated by the hospital) to be trained in abroad for a period of 2 weeks to familiarize with advanced different special applications — including all costs.(If Applicable)
23 | Test Equipment & Phantom required for quality assurance & performance
24 | Interior & lighting ambience for patient relaxation environment
25 | All accessories to improve Work flow, enhance and speed the centring and positioning of patient
26 | Consumables
26.1 | Supplied with all standard consumables.
27 | Available Additional Option
27.1 | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
28 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Radiology technicians (End User Training) and doctors is required.
3 | Supplier must submit FOC technical training program for Biomedical Engineer including all expenses
4 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
5 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
6 | A Price List For Spare Parts: Must be provided to cover at least 10 years of operation and maintenance of the equipment supplied.
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
10 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
11 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
12 | Warranty must be at least 5 years comprehensive from date of installation and equipment acceptance.
13 | Vendor Must include Site preparation (pre-Installation) work for the proposed model

## Civil & MEP Works
Civil and MEP works for CT Scan Unit
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Civil and MEP works for CT Scan Unit | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
All Pre-installation requirements should be submitted, including suggested room Design and dimensions
1 | All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required.
2 | Scope of Proposed Design and As built Drawings ( after completion of Project)Should be Included
3 | Interior Designing and Decorative works to be Proposed and should Include in scope of Works
4 | Demolition and Mobilization works should be included if required
5 | Mobilization(Wall opening and closing ) for Imaging machine delivery should be included
6 | The system offered shall be designed to operate normally under the conditions of UAE, Environmental factors, FNAR and MOH policy the conditions include Power Supply,Climate,Temperature, Humidity,Dust,etc
7 | Detailed Lead test reports for existing lead and compliance certification to be provided at each step to ensure the radiation leakage not only meets but also exceeds the manufacturers requirements.
8 | Provide overlap of the material at each panel joint for the maximum level of performance, attenuation integrity and electrical conductivity.
9 | To install appropriate internal fit out or lining, lighting and electrical meet international standards.
10 | supply and installation of Warning light and signs
11 | Partition works and Ceiling works should be Similar to UHS standards. Proposed Specification need to be provided
12 | Aesthetic Ceiling picture can be done in X-ray room as per client suggestion
13 | Acoustic ceiling tiles.
14 | The Sound insulation can be incorporated into the framework of the walls during assembly and to ensure standard acoustic internal lining.
15 | Wall treatment
16 | To provide clearest observation window with observation glass
17 | The observation Glass shall be anti-reflective coating
18 | Flooring- Medical Grade (similar or better than existing flooring at UHS)
19 | Removing of existing flooring & levelling with self-levelling screed for complete area.
20 | Power Supply Capacity and Specifications
21 | Proposal for Supply and Installation of Electrical works. socket and switches should be via brand and of same type used in UHS premises.(All wires will be concealed and in conduct pipes and the wiring for electrical and network cable will be commensurate with the existing ).
22 | The lighting should be all LED and should support energy conservation
23 | HVAC- Study the current layout and Propose the modifications required
24 | Fire Alarm and Fire Fighting modification works to be propose
25 | Supply and installation Bedhead units for proposed locations including Medical Gas Points and Nurse call system
26 | Existing ceiling mounted procedure light and ceiling mounted lead shield can be used.
27 | All MEP work including medical gas system should be according to UHS stander which included future line if needed and current system assessments to avoid any struggling with operation after the installation
28 | lead lining should be according to the FANR requirements As per the safety assessment ,Tube outcome ,Generator power and machine orientation.
29 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
30 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Equipment. Two Copies each and one Soft Copy for each.
3 | It is mandatory for the Contractor to submit End-of-Sale, End-Of-Support, and End-Of-Life for each individual hardware component - Proof documents from the manufacturer to be attached with the proposal.
4 | Proposal Should be submitted after having site survey and comparison with the current As built
5 | All The works Should Comply with Standards of Civil defense, SEWA and MOH
6 | Contractor should discuss the final technical proposal with the technical team before submission. Proposal should include ongoing warranty, support.