## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Ultrasound Machine for OBG- OPD | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Ultrasound machine and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | Exam Types Abdominal, obstetrical, gynecological
3 | Operating Modes B-mode, M-mode, PW mode ,CW mode, Color doppler, Power doppler mode, Volume Mode (3D/4D), Color M-mode, HD live, Volume contrast Imaging (VCI), Advanced Volume Imaging (AVC, VCAD, VOCAL), STIC, Tissue Doppler Imaging (TDI) , THI, TUI , SRI, V-SRI , CE.
4 | Provides real-time volumetric imaging (3D/4D) for detailed fetal assessments
5 | Volume Navigation
6 | Four active probe slots to facilitate quick switching between transducers including: Three Transducers - 1)2D Transvaginal Probe 2)3D Transabdominal Probe 3)4D Transabdominal Probe
7 | •Monitor Screen minimum 19-23′′ High-Resolution LCD •Min 10′′High Resolution Touch LCD screen for intuitive operation
8 | Specify Dimensions and weight of the unit
9 | Specify Min/Max Hight adjustment
10 | Unit should facilitates seamless integration with PACS system
11 | Peripheral Options • User Interface Operator Keyboard • Digital B&W thermal printer • DVD video recorder • Foot Switch with programmable functionality •Gel warmer
12 | Screen Display Patient name, patient ID, Hospital name, GA or LMP, Birth date, date and time, probe name, application name, mode, gray scale bar, depth scale, focal zone marker, frame rate
13 | Measurements Generic measurements, Abdomen calculations, obstetrics calculations, obstetrics fetal echo, gynecology calculations, pediatric calculations, OB tables, fetal weight estimation, fetal ratios
14 | Image Storage •Storage Formats:DICOM – compressed/uncompressed, single/multiframe with/ without RawData • Hard Disk capacity for both image and raw data should be 2 TB •Export JPEG, WMV (MPEG 4), and AVI formats •Support Storage Devices:USB Memory Stick,CD,DVD, Hard Drive storage.
15 | Unit should be provided with all necessary accessories and consumables.
16 | Connectivity & DICOM •Ethernet network connection• Verify• Print• Store• Modality Worklist• Storage Commitment• Query / Retrieve
17 | Full DICOM and HL7 compatibility
18 | Frame rate/Sampling rate must not be less then 150 frames/sec
19 | Unit Must have HDMI /DVI/VGA ports to connect with extenal patient view monitor
20 | Must have universal power adapter 220-240v ,50/60HZ
21 | Minimum One Terra Byte Internal Hard Disk
22 | Equipped with an internal battery for emergency shutdowns
23 | Quality and Patient safety
24 | Unit must be FDA / CE marked .
25 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
26 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
27 | Accessories
Should have All standard Accessories must be included as per manufacturer .
28 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
2
3
29 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training if or operator and Physians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation