## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| CSSD Washer Disinfector with 18 DIN Trays | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1.1 | Specify Make/Model/Origin
2 | Purpose: The washer disinfector is required for the cleaning, disinfection, and drying of surgical instruments, medical devices, and other hospital tools in compliance with infection control standards.
3 | Single chamber, double door washer disinfector compliance with EN / ISO 15883 certified.
4 | Tray Capacity: Washer disinfector capacity of 18DIN trays with adjustable rack options to ensure proper placement of various surgical instruments.
5 | Chemical Disinfection: Capable of using chemical disinfectants or detergents for disinfection as per manufacturer recommendations.
6 | Cycle Time: The total cycle time for cleaning, disinfection, and drying should be specified, preferably within the range of 30–60 minutes for efficiency. The unit should support both low and high-level disinfection processes.
7 | Washer should have LED touch-screen control panel
8 | Washer should have chamber Illuminated with long life illuminating LED light
9 | Control System: Fully automatic, microprocessor-based with touchscreen display
10 | Washer should have Audible and visual warning system
11 | Must be equipped with Temperature Detectors to sense temp inside chamber
12 | Detergent Dosing: Automatic dosing for detergent, neutralizer, and rinse aid
13 | Must be equipped with chemical dose Detectors to sense chemical doze inside chamber
14 | Safety Features: Door interlocks, over-temperature protection, emergency stop
15 | Must be equipped with spray wings between each shelf to allow water to reach all load surfaces
16 | Must have a water level sensor in the wash chamber sump
17 | Washer disinfector should be compactible with Robotic instruments washing.
18 | To be equipped with drying system with HEPA filter providing entire chamber coverage
19 | Must provide Built-in chemicals cabinet
20 | Programs: Minimum 10 to 12 programmable cycles (pre-wash, Main wash, rinsing, thermal disinfection(high-level), drying)
21 | Emergency stop switch @ Loading & Unloading side
22 | Doors equipped with obstruction sensors
23 | Safety Lockout feature to abort cycle when door is opened
24 | Door interlock feature to prevent cross-contamination
25 | Washer disinfector should be environmental friendly, Water saving ,Energy saving etc..
26 | The washer disinfector must allow the adjustment of temperature for various phases of the cleaning process, with a range of 55°C - 95°C.
27 | Cycle Documentation: Inbuilt printer and connected to the computer (for Instrument tracking system ).
28 | Include racks and inserts for surgical instruments, MIS instruments, and anesthesia items.
29 | Vendor must ensure proper installation, validation (IQ/OQ), and performance qualification (PQ).
30 | Operator training and maintenance manual must be provided.
31 | The supplier shall carry out necessary stainless steel cladding work around the washer disinfector to ensure effective separation between the decontamination area and clean area. The cladding must be made of high quality corrosion resistant medical grade stainless steel.
32 | The unit should provide temperature validation to ensure all cycles meet the required sterilization standards, and Environmental temperature :
33 | Dimensions: Specify dimension of the unit Provide the external and internal dimensions of the unit to ensure compatibility with the available installation space.
34 | Unit must be Capable to add any upcoming new upgrades. (Software or Hardware)
35 | weight of the unit
36 | Quality and Patient safety
36.1 | Unit must be FDA / CE certified
36.2 | Unit must comply ,IEC 60601-1 ; ANSI S3.39 ; ISO 60645, ISO13485,EN/ISO 15883 Quality management system for medical device or equivalent .
36.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
37 | Accessories
Supplied with all standard accessories.
38 | Consumables
38.1 | Supplied with all standard consumables.
39 | Available Additional Option
No. | OPTIONS DESCRIPTION
40 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for and for CSSD technicians and Physicians (End User Training) if required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .
I. | Vendor Must include Site preparation (pre-Installation) work for the proposed model