## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| New born Hearing Screener ( DIAGNOSTIC UNIT) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | specify Make/Model/Origin
3 | Device must be with latest technology and standard of care and User-Friendly Interface
4 | Machine must be available to do Several test like; Automated Auditory Brainstem Response (AABR),Automated otoacoustic emission(AOAE),Distortion product otoacoustic emission (DPOAE),Transitory evoked otoacoustic emission (TEOAE) Screening.
5 | Accuracy:AABR testing must be accurate to within ±5 dB of recognized standards.OAE testing must meet accuracy standards as per the relevant clinical guidelines.
6 | Unit must able to detect Noise during testing
7 | specify test result AABR/OAE timing in sec/minutes
8 | Frequency Range:AABR: 500Hz to 4000Hz and OAE: 1000Hz to 4000Hz.
9 | Sensitivity:1)ABR Sensitivity: Must detect hearing impairment as low as 30-40 dB HL (Hearing Level)
10 | 2)OAE Sensitivity: Must detect outer hair cell function with a threshold of 30-40 dB HL.
11 | Test Pass Criteria: The system should display a clear "Pass" or "Refer" outcome based on standardized criteria.
12 | Must be equipped with Color touch Screen , External monitor option, Printer and other accessories .
13 | Connectvity:1)Data Management: Integration with hospital information systems (HIS) 2)Data Export: Ability to export results in PDF 3)Wireless Connectivity: Wi-Fi, Bluetooth, or USB for transferring data to external systems
14 | Battery: Rechargeable lithium-ion battery for portable use (minimum 2-4 hours of continuous operation).
15 | Environmental temperature
16 | Specify dimension of the unit
17 | weight of the unit
18 | Quality and Patient safety
18.1 | Unit must be FDA / CE certified
18.2 | Unit must comply ,IEC 60601-1 ; ANSI S3.39 ; ISO 60645, ISO13485 Quality management system for medical device or equivalent .
18.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
19 | Accessories
Supplied with all standard accessories.
20 | Consumables
20.1 | Supplied with all standard consumables.
21 | Available Additional Option
No. | OPTIONS DESCRIPTION
22
Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for and for lab technicians and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .