## Indirect Ophthalmoscope, etc
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Indirect Ophthalmoscope | 2 | Comply | Supporting Documents Ref. | by Hospital
2 | Direct Ophthalmoscope | 5
3 | Portable Slit Lamp | 3
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Indirect Ophthalmoscope
1.1 | Please specify make, model, and origin of the machine. |  | Yes
1.2 | Device must have LED light source with continuously adjustable illumination intensity with a maximum of no less than 500 Lx
1.3 | The design of the device must be lightweight and allowing for easy and convenient operation
1.4 | Equipped with portable power supply
1.5 | Device is equipped with color filter,teaching mirror and sclera depressor
1.6 | Device must have adjustable headband to accommodate for different patient sizes
1.7 | Please specify pupillary adjustment (52~74 mm)
1.8 | Please specify illumination area
1.9 | Must have big, middle, and small light spots
1.1 | Must have red-free, white and cobalt blue filters
1.11 | Please Specify working distance
2 | Direct Ophthalmoscope
2.1 | Please specify make, model, and origin of the machine.
2.2 | Device must be light weight and easy to use
2.3 | Device must use an LED light source with 3000K color temperature
2.4 | Device must have pupillometer gauge for easy estimation of pupil size
2.5 | Device must have red-free, green, and cobalt blue filter
2.6 | At least +29D to -30D dioptre range
2.7 | Device must have changeable apertures
2.8 | Device must have adjustable magnifications
2.9 | Device must be able to withstand hospital cleaning and disinfection processes
3 | Portable Slit Lamp
3.1 | Please specify make, model, and origin of the machine.
3.2 | Slit length of 12mm
3.3 | LED Light source
3.4 | Magnification modes of 10x and 16x
3.5 | Please specify working distance
3.6 | 13° Angle of convergence
3.7 | Field of view of at least 16mm with 10x magnification and 10 mm with 16x magnification
4 | Quality and Patient Safety
4.1 | FDA approved / CE Marked
4.2 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
4.3 | Offered System must comply with Electrical safety IEC 6060-1
4.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Autorefractometer With Air Puff
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | All in one Autokerato Refractometer, tonometer and pachymetry | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
| Please specify make, model, and origin of the machine. |  | Yes
A | Device must contain space saving autoref , kerato , noncontact tonometer and pachymeter
B | Large pupil zone imaging method
C | Optimal fogging to minimize accommodation
D | Automated calculation of corrected IOP
1 | Auto refractometer
Measurement range :
1.1 | Auto keratometer ( Please specify the in built technology)
1.2 | Please specify the Measurement range
1.3 | Curvature radius 5.00 to 13.00 mm (0.01 mm increments)
1.4 | Refractive power 25.96 to 67.50 D (0.12 / 0.25 D increments)
1.5 | Cylindrical power 0 to 12.00 D (0.12 / 0.25 D increments)
1.6 | Axis 0 to 180º (1° / 5° increments)
1.7 | Pupil size measurement range (Angle ,2.0mm)
1.8 | 1.0 to 10.0 mm (0.1 mm increments)
1.9 | PD measurement range ( 1mm , 20 to 85mm)
1.91 | Must be having auto fog system
2 | NON-CONTACT TONOMETER :
2.1 | Measurement range
2.2 | 1 to 60 mmHg (1 mmHg increments)
2.3 | Please specify the Working distance
3 | NON CONTACT PACHYMETER
3.1 | Please specify the Measurement range
3.2 | 400 to 750 µm (1 µm increments)
4 | Printer
4.1 | Preferably : Thermal line printer with easy loading and auto cutter
5 | Chinrest travel distance :
5.1 | Up/down : 65- 70mm
6 | Quality and Patient Safety :
6.1 | FDA approved / CE Marked
6.2 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
6.3 | offered System must comply with Electrical safety IEC 6060-1
6.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Trial Lens Set,etc
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Trial Lens Set | 1 | Comply | Supporting Documents Ref. | by Hospital
2 | Ishihara Color Vision Chart (Book) | 1
3 | Prism Lens Tools | 1
4 | Stereo Acuity Chart (Book) | 1
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Trial Lens Set
1.1 | Please specify make, model, and origin of the machine.
1.2 | Trial frame should be made of metal
1.3 | Trial lens set should be minimum 167 pcs
1.4 | Trial lens set should have spherical, cylindrical, and prismatic lenses that have a minimum range of (Spherical: +/- 0.25 to +/- 16.00D) (Cylindrical: +/- 0.25 to +/- 4.00D) (Prismatic: 1 to 5D)
1.5 | Trial set and frame should come in a compact carrying case
2 | Ishihara Color Vision Chart (Book)
2.1 | Book should contain all the standard and appropriate designs used to test for color blindness.
3 | Prism Lens Tools
3.1 | Prism Lens tools
4 | Stereo Acuity Chart (Book)
4.1 | Chart or book should contain all appropriate designs necessary for testing stereoscopic acuity.
5 | Quality and Patient Safety
5.1 | FDA approved / CE Marked
5.2 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
5.3 | Offered System must comply with Electrical safety IEC 6060-1
5.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Auto Lens Meter, etc
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Auto Lens Meter | 1 | Comply | Supporting Documents Ref. | by Hospital
2 | Retinoscope | 2
3 | Tonopen | 3
4 | Digital Visual Acuity Chart | 2
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Auto Lens Meter
1.1 | Please specify make, model, and origin of the machine. |  | Yes
1.2 | Device should accurately and automatically measure the prescription of the lens
1.3 | Hartmann sensor with at least 108 measurement points
1.4 | Scale mode function
1.5 | UV transmittance measurement
1.6 | Vertical 5.7-inch color LCD touch screen or larger
1.7 | Sphere lenses: 0~±25D or more
1.8 | Cylinder lenses: 0~±10D or more
1.9 | Cylinder Axis Angle: 0°~180° or more
2 | Retinoscope
2.1 | Please specify make, model, and origin of the machine.
2.2 | Device should be lightweight and easy to handle
2.3 | Device should accurately and quickly measure the astigmatism axis.
2.4 | Filament rotation of 360 degrees with upward and downward movement
2.5 | The brightness of the streak can be adjusted
2.6 | Streak rotation of 360 degrees
3 | Tonopen
3.1 | Please specify make, model, and origin of the machine.
3.2 | Device must be lightweight and easy to use with an ergonomic design
3.3 | Device must be able to provide quick and accurate emasurements of IOP
3.4 | Device must not require routinely calibration
3.5 | Device must be gravity independent allowing it to be used in any position
3.6 | Device must have a measurement range of at least 5-55 mmHg
4 | Digital Visual Acuity Chart
4.1 | Please specify make, model, and origin of the machine.
4.2 | Device remote must have luminous keypads to be easily usable in the dark
4.3 | Omnidirectional remote control
4.4 | Remote control should have a long working distance
4.5 | HDMI and VGA supported
4.6 | Compact size
4.7 | Device should be easy to set up
4.8 | Testing distance from at least 1m to 7m
4.9 | Random sequencing of optotypes
4.1 | Should fit any monitor or TV set
4.11 | High definition, up to 1920 x 1080
5 | Quality and Patient Safety
5.1 | FDA approved / CE Marked
5.2 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
5.3 | Offered System must comply with Electrical safety IEC 6060-1
5.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Wide Field Fundus Camera
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Wide Field Fundus Camera | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
| Please specify make, model, and origin of the machine. |  | Yes
1 | Must provide 180 degrees field of view of the retina
2 | Image must be high resolution
3 | Device must use an LED light source
4 | Device must include multiple imaging modes including true color (with red, green, blue channel seperation) and infrared
5 | Minimum pupil diameter of at least 2.5 mm
6 | Please specify working distance
7 | Device must be wheelchair accessible
8 | FDA approved / CE Marked
9 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
10 | Offered System must comply with Electrical safety IEC 6060-1
11 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Opthalmic Unit and Dr Chair
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Opthalmic Unit with Patient Chair | 1 | Comply | Supporting Documents Ref. | by Hospital
2 | Revolving Examiner's Chair | 1
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Patient Chair for Opthalmic Examination
1.1 | Unit must include a sliding tabletop that can accommodate at least two instruments
1.2 | Please specify make, model, and origin of the machine.
1.3 | Chair should be electrically controlled
1.4 | Chair should be able to move up and down with a minimum range of 568~708mm
1.5 | Chair should be rotatable 0°~270°
1.6 | Backrest adjustable angle 80°~180°
1.7 | Headrest should be adjustable
1.8 | Chair max load at least 150 Kg
2 | Revolving Examiner's Chair
2.1 | Please specify make, model, and origin of the machine.
2.2 | Height of the chair should be easily adjustable
2.3 | Chair should be rotatable 360 degrees smoothly
3 | Quality and Patient Safety
3.1 | Device should be CE and FDA approved
3.2 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
3.3 | Offered System must comply with Electrical safety IEC 6060-1
3.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Tear Osmolarity Test Machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Tear Osmolarity Test Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
| Please specify make, model, and origin of the machine. |  | Yes
1 | Device must quickly and accurately measure tear osmolarity to aid in the diagnosis of dry eye disease
2 | Device must be lightweight, ergonomic, and easy to use repeatedly
3 | LCD Display
4 | Measurement range of at least 275–400 mOsms/L
5 | Device has internal battery
6 | Device uses single-use sensors
7 | Device should be CE and FDA approved
8 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
9 | Offered System must comply with Electrical safety IEC 6060-1
10 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .