## UBT
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Fully automated UBT(Urea Breath Test) & Hydrogen Breath Test for SIBO Analyzer system |  | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the fully automated UBT analyzer with User-friendly interface & Flexible data management options.Please specify the details clearly and all the requested specs information must be well supported with the reference documents.
2 | Fully automated analyser should incldue UBT & SIBO analysis preferably both testing in a same analsyer or separate unit | 1
Detection of Initial infection and confirmation of eradication
Adult & Paediatric testing should be done with out manual calculation
Quality control and Consumable should have long expiry date
3 | Pre-Analytical
Easy/simple patient preparation method prior to sample collection
Minimum interference of Antibiotics & Drugs
Easy sample collection method with minimum intervention of time
UBT Bag identification needed for Pre and post samples
Sample stability should lasts for 1-2 weeks
4 | Analytical
Easy loading and unloading of samples with Minimum Carryover
QC, Calibration,Maintanance & Troubleshooting procedure must be user friendly
Minimum 15 tests per hour with reduced processing time
Minimum 3-4 sample should be done at a time
5 | Post Analytical
Interface and data management options- includes Barcode scanner, print, export, or transmit results
Should be compatible with Intersystem LIS
Availablity of sufficient technical support /Engineer onsite to support for breakdown
6 | Quality requirements
FDA/CE mark approval
Number of customers used in CAP/ISO accredited labs in UAE / ME Region
External quality control program enrollment
Validation of analyser should be throughly supported by vendor along with providing supporting documents
Availablity for demo
7 | Sample Consumption
UBT - 25-40 Samples per week , SIBO <10 samples per week
8 | Accessories
Supplied with all standard accessories.
9 | Consumables
Supplied with all standard consumables.
10 | Available Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for lab staff and later as needed .
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS,CAP
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warrenty : Comprehensive minimum 5 Years warrenty .

## Automated MI & ASP
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Automated Microbial Identification and Antibiotic sensitivity Platform | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the fully automated compact solution for fast, accurate microbial ID/AST testing
2 | Fully automated microbiology workflow system should include Maldi Tof ID & AST (Preferably in a single system)
2.1 | Maldi Tof system for rapid and accurate identification of bacteria, mycobacteria and fungal pathogens | 1
Throughput of 100 tests/run
Minimal sample preparation steps with auto dispensing feature
Extensive identification menu
Features to highlight unusual phenotype
Calibration Frequency?
Easy Maintenance & troubleshooting
Extended reagent stability & supply should be smooth to have minimal Lot to Lot variations.
2.2 | Automated Antibiotic sensitivity Platform | 1
Throughput of 100 tests/day
Reduce time to microbial identification and antibiotic susceptibility testing results
Extensive susceptibility menu
Extended MIC range to enable low-level resistance detection
Calibration Frequency?
Easy Maintenance & troubleshooting
Extended reagent stability & supply should be smooth to have minimal Lot to Lot variations.
2.3 | Middle ware compatable to integrate Maldi Tof and AST systems | 1
System should have integrated powerful middle ware with ablitiy of automatically providing validation of every susceptibility test result
Optimize laboratory efficiency on identification & senstivity reporting
LIS-compatible software that evolves with microbiology updates.
Ability to accomodate Bacteriology Taxonomical changes
Should Support CLSI M100 updates.
Designed to provide ID/AST results in shortest TAT.
Automatic validation & transfer of preliminary results stopping transmission of discrepant reults to LIS unless reviewed
Easy-to-access results with multiple filters
Bidirectional connection with LIS- ust be compatiblw with Intersystem TCLE system
3 | Environmental temperature and heat dissipiation from the unit in BTU
4 | Specify dimension of the unit
5 | weight of the unit
6 | Quality requirements
6.1 | FDA/CE mark approval
6.2 | Number of customers used in CAP accredited labs in UAE / ME Region
6.3 | CAP external quality control program enrollment
6.4 | Availablity for demo, ability to provide back up during break down.
7 | Sample Consumption
Number of isolates performed - 60 /day
8 | Accessories
Supplied with all standard accessories.
9 | Consumables
Supplied with all standard consumables.
10 | Available Additional Option
No. | OPTIONS DESCRIPTION
11 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for lab staff and later as needed .
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS,CAP
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warrenty : Comprehensive minimum 5 Years warrenty .

## POCT- ACT
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | POCT ACT (Activated Clotting time) ANALYZER | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Specify Make/Model/Origin
3 | Size: Small/ handheld device.
4 | Easy to use
5 | Connection to track care
6 | Software: Ability to view the device from a software or by the connection of a middle ware (AQURE).
7 | Maintenance: Auto-maintenance.
8 | Operator Lock-out (Multiple users)
9 | Calibration due lockout (if needed)
10 | Automatic Quality control lockout
11 | Automatic Electronic Quality control (every 8 hours)
12 | Certification: FDA or CE certified
13 | Quality control: Manufacturer Quality control availability
14 | External Quality control: Listed in CAP PT
15 | Sample type: Finger prick (Availability of capillary tube with the kit)
16 | Sample volume: 100 microlite or less
17 | Automatic barcode identification of samples
18 | User friendly, Easy loading and unloading of samples
19 | Quality and Patient safety
20 | Unit must be FDA / CE certified
21 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
22 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
23 | User-friendly operation and smooth movement - One-unit package with in-built battery/generator system - Durable with operational implication for different position and parts of the body.
22 | Integration
The Point-of-Care Testing (POCT) devices must be integrated with the Laboratory Information System (LIS) to enable unidirectional transmission of test results. This integration should use the HL7 v2.6 standard, specifically the ORU^R01 message type, to facilitate the transfer of observation results from POCT devices to the LIS. The system must support patient identification through HL7-ADT message and barcode scanning both to ensure accurate and efficient patient data capture at the point of care. It should also provide secure data transfer, robust error-handling, and acknowledgment mechanisms to guarantee reliable and accurate result reporting from POCT devices to the LIS in real-time, while ensuring data integrity and compliance with industry standards.
23 | Accessories
23.1 | Supplied with all standard accessories.
24 | Consumables
24.1 | Supplied with all standard consumables.
24.2 | Monthly usage is 50T per month
25 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
24 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at least 3 years comprehensive from date of installation and equipment acceptance.