## Glucometer
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | POCT Glucometer Analyzer (40 device with docking station) | Sys | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Specify Make/Model/Origin
3 | Size: Small/ handheld device.
4 | Usage: Easy to use
5 | Tests
6 | 1- Glucose - Long stability & expiry of strips and solution
7 | 2- Ketone - Long stability & expiry of strips and solution
8 | Built in scanner (Preferred to read 1D & 2D Barcode)
9 | Software: software remote access (AQURE is prefered)
10 | Maintenance: No-maintenance.
11 | Custamble Operator Lock-out (Multiple users).
12 | Access expiry date & Lot number.
13 | Provide L-J Charts and statistics for QC through the software
14 | Automatic Quality control lockout
15 | Test Comment option & Meal selection (Fasting, Random etc.)
16 | Certification: FDA or CE certified
17 | Test duration: < 20 Seconds
18 | Quality control: Quality control provided by Manufacturer
19 | External Quality control: Listed in CAP PT
20 | Sample type: Finger prick.
21 | Sample volume: 100 microlite or less
22 | Other installation based in UAE
23 | Quality and Patient safety
23.1 | Unit must be FDA / CE certified
23.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
23.4 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
24 | Waived Glucose and ketone
Integration
25 | The Point-of-Care Testing (POCT) devices must be integrated with the Laboratory Information System (LIS) to enable unidirectional transmission of test results. This integration should use the HL7 v2.6 standard, specifically the ORU^R01 message type, to facilitate the transfer of observation results from POCT devices to the LIS. The system must support patient identification through barcode scanning to ensure accurate and efficient patient data capture at the point of care. It should also provide secure data transfer, robust error-handling, and acknowledgment mechanisms to guarantee reliable and accurate result reporting from POCT devices to the LIS in real-time, while ensuring data integrity and compliance with industry standards.
26 | Accessories
26.1 | Supplied with all standard accessories.
26.2 | 40 device with docking station
27 | Consumables
27.1 | Supplied with all standard consumables.
27.2 | Total usage of average 10,000 strip per month
28 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
28 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Clinical Staff (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at least 3 years comprehensive from date of installation and equipment acceptance.
12 | Provide Devices based on consumble (Rental Based).

## POCT ACT
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | POCT ACT (Activated Clotting time) ANALYZER | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Specify Make/Model/Origin
3 | Size: Small/ handheld device.
4 | Easy to use
5 | Connection to track care
6 | Software: Ability to view the device from a software or by the connection of a middle ware (AQURE).
7 | Maintenance: Auto-maintenance.
8 | Operator Lock-out (Multiple users)
9 | Calibration due lockout (if needed)
10 | Automatic Quality control lockout
11 | Automatic Electronic Quality control (every 8 hours)
12 | Certification: FDA or CE certified
13 | Quality control: Manufacturer Quality control availability
14 | External Quality control: Listed in CAP PT
15 | Sample type: Finger prick (Availability of capillary tube with the kit)
16 | Sample volume: 100 microlite or less
17 | Automatic barcode identification of samples
18 | User friendly, Easy loading and unloading of samples
19 | Quality and Patient safety
20 | Unit must be FDA / CE certified
21 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
22 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
23 | User-friendly operation and smooth movement - One-unit package with in-built battery/generator system - Durable with operational implication for different position and parts of the body.
Integration
24 | The Point-of-Care Testing (POCT) devices must be integrated with the Laboratory Information System (LIS) to enable unidirectional transmission of test results. This integration should use the HL7 v2.6 standard, specifically the ORU^R01 message type, to facilitate the transfer of observation results from POCT devices to the LIS. The system must support patient identification through barcode scanning to ensure accurate and efficient patient data capture at the point of care. It should also provide secure data transfer, robust error-handling, and acknowledgment mechanisms to guarantee reliable and accurate result reporting from POCT devices to the LIS in real-time, while ensuring data integrity and compliance with industry standards.
25 | Accessories
25.1 | Supplied with all standard accessories.
26 | Consumables
26.1 | Supplied with all standard consumables.
27 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
24 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and Clinical Staff (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at least 3 years comprehensive from date of installation and equipment acceptance.