## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Surgical microscope for Neuro & Spine surgeries | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | System must be state of the art , latest release model , having Ergonomic design and easy to use not limited to sharp images, flexibility, optimum maneuverability, connectivity with other devices in the OR and high-quality video and recording systems.
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
3 | Additional 24 to 26″ Full HD Widescreen Color Monitor displays a bright and clear image of the surgeon’s view in real-time for teaching and educational purposes.
4 | Multi-Specialty Application: Brain aneurysm repair, Tumor resections, Arteriovenous malformation (AVM) treatment, Cerebral artery bypass surgery, Epilepsy surgery,Indocyanine green (ICG) angiography,5-ALA (vascular and brain tumors) & Spine surgery.
5 | Specify Unit Make ,Model and country of origin
6 | Apochromatic lens:
6.1 | Motorized focus working distance 200 – 625 mm or more
6.2 | Motorized zoom; zoom ratio 1:6,magnification factor
6.3 | Magnetic wide field eyepieces
7 | Illumination:
7.1 | 2 x 300 W Xenon, with automatic lamp exchange
7.2 | Automatic Iris Control for adjusting the illumination to the field of view
7.3 | Focus Light Link: working distance controlled light intensity
7.4 | Individual light threshold setting
7.5 | Additional illumination beam to brighten up shadows, motorized
8 | System Operation:
8.1 | Multifunctional programmable handgrips
8.2 | Magnetic clutches for all system axes
8.3 | Central user interface with full-screen video
8.4 | variable speed & active damping
8.5 | Manual and motorized Point Lock functions
9 | System setup and operation:
9.1 | Auto Balance
9.2 | Park & drap Position
9.3 | Direction of View 45° upwards
9.4 | Shaft Diameter 3.6 mm Shaft Length 120.0 ± 1.0 mm
9.5 | Total Diameter 13.0 mm Field of View 100° +5°/-10°
9.6 | Illumination 20 – 35 lumen LED
9.7 | Good Image Resolution 1920 × 1080 pixel or higher
9.8 | Operation Temperature +10 to +40 °C
10 | Video recording & Storage :
10.1 | Integrated Full HD video camera- 1080p
10.2 | 24" HD video touchscreen with extendable arm to control all the features of the recorder. Includes secondary monitor 4K as optional
10.3 | Integrated still image capturing both on HDD and USB-media
10.4 | Large HDD support long recording time internal hard disc 2 to 4TB capacity or Expandable
10.5 | Video recording integration with Hospital Information system (HIS)
10.6 | Unit must be HL7 compatible
10.7 | Integration should facilitate workflow management and enable access to worklists and storage from the HIS.
10.7 | Required Hardware(Server & storage )should be included in the proposal as optional
11 | Unit must be antimicrobial Protection
12 | Specify Dimensions and weight of the units
13 | Unit should be provided with all necessary accessories and consumables.
14 | Must have universal power adapter 220-240v ,50/60HZ
15 | Specify the power consumption for the Offred units and heat dissipation(BTU/hr)
16 | Accessories:
16.1 | Functions, Wired or wireless foot control
16.2 | Magnetic wide field eyepieces with integrated eyecups
16.3 | 3-step magnification changer
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ,ISO8600, ISO13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
18 | Accessories
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Available Additional Option
No. | OPTIONS DESCRIPTION
21 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End user / application training must be done for Endoscopy doctors and Nurses.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .