## Ultrasound
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| PORTABLE VASCULAR ULTRASOUND FOR OPD | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Ultrasound machine and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | Premium ultrasound machine capable of top-of-the-line and 3D imaging along with clear imaging for Vascular department. Ultrasound machine is suitable for accurately assessing vascular conditions, including conditions like deep vein thrombosis (DVT), atherosclerosis, carotid artery stenosis, and other vascular abnormalities.
3 | Higher-end Ultrasound machine which includes all Vascular application not limited to Abdominal, Vascular /Peripheral, Carotid Artery Imaging, peripheral Arterial Disease (PAD) Screening, Deep Vein Thrombosis (DVT) Detection, Chronic Venous Insufficiency (CVI),Aortic Aneurysm Surveillance, Varicose Vein Assessment, Arteriovenous Malformations (AVMs),Vascular Tumour Detection, Arterial Stenosis and Plaque Evaluation. |  |  |  |  | 3
4 | Should have two to four Active Probe slots.
5 | Transducer Types: • Convex probe • Linear array probe • Phased array probe •High-frequency probes (7-15 MHz) •Lower-frequency probes (2-5 MHz)
6 | Volume Navigation
7 | Must include Operating Modes • B-Mode •Colour Doppler: Anatomical & Colour M-Mode • Colour Flow Mode (CFM) • Extended Field of View • Power Doppler Imaging (PDI) • Pulse Wave Doppler • CW Doppler • Tissue Doppler• •Power Doppler •Spectral Doppler •3D •4D(optional) • Colour Doppler • Compound Imaging • Contrast (optional) • DICOM, •Elastography •HD-Flow •HPRF • M-Mode Flow • Needle Visualization -(as optional item) • SonoVCAD, •Speckle Reduction •STIC • • Tissue Harmonics •VOCAL •Volume Speckle Reduction •Coded Excitation/Harmonics
8 | Unit should be fully supported to PACS system-Capability to store and archive images and video clips in various formats (e.g.DICOM, JPEG)
9 | •Monitor Screen minimum 19-21′′ High-Resolution LCD •Min 10′′High Resolution Touch LCD screen for software menu
10 | specify Min/Max Hight adjustment
11 | Horizontal/Vertical viewing angle of monitor
12 | Image Storage: •Storage Formats:DICOM – compressed/uncompressed, single/multiframe with/ without RawData •Export JPEG, WMV (MPEG 4), and AVI formats • support Storage Devices:USB Memory Stick,CD,DVD, Hard Drive storage
13 | Peripheral Options • User Interface Operator Keyboard • Digital B&W thermal printer • DVD video recorder -(as optional item) • Digital color thermal printer -(as optional item) • Foot Switch with programmable functionality •Gel warmer
14 | Connectivity & DICOM : • Ethernet network connection• Verify• Print• Store• Modality Worklist • Storage Commitment• Query / Retrieve
15 | Having B/M color and PW/CW
16 | Unit Must have HDMI /DVI/VGA ports
17 | Frame rate/Sampling rate must not be less then 150 frames/sec
18 | Must be supplied with 1 TB HDD or more
19 | Full DICOM and HL7 compatibility
20 | Specify Dimensions and weight of the unit
21 | Unit should be provided with all necessary accessories and consumables.
22 | Must have universal power adapter 220-240v ,50/60HZ
23 | Quality and Patient safety
24 | Unit must be FDA / CE marked .
25 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
26 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1.
27 | Accessories
Supplied with all standard accessories.
28 | Consumables
Supplied with all standard consumables.
29 | Available Additional Option
No. | OPTIONS DESCRIPTION
30 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .