## Ultrasound
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Central Monitoring System ( 4 Workstations + 1 Main view station) | 5 | Comply | Supporting Documents Ref. | by Hospital
2 | Cardiac Monitor | 24
| Specifications for Cardiac Monitor |  | Yes/No
3 | Technical offer must be clear in terms of standards , which will be offered with the central monitor with cardiac monitors. Items should be mention clearly. Please specify the details clearly and all the requested specs information must be well supported with reference documents.
4 | Specify Make, Mode and Origin of the product.
5 | Touchscreen Display size 12 inches with good resolution
6 | 5 ECG channels and Respiration rate monitoring
7 | SpO2 Monitoring
8 | NIBP Monitoring with all sizes of cuffs (Adult, Large adult- Two sets)
9 | Temperature Monitroing
10 | Back-up Li-ion battery up to 2 hours operation for transport scenarios
11 | Visual and audible color coded alarms with high or low values
12 | Patient data management
13 | Unit must has antimicrobial protection
14 | Save patient data to the CMS archive
15 | Full DICOM and HL7/eGateway compatibility , Integrated with HIS
16 | Can be integrated with the Central Monitoring System
17 | Specify Dimensions and weight of the unit
18 | Unit should be provided with all necessary accessories and consumables.
19 | Must have universal power adapter 220-240v ,50/60HZ
| Specifications for Central Monitoring System |  | Yes/No
20 | Can be Integrated with the given monitors
21 | Multi-bed patient data visualization - Workstation with 25 Bed licenses
22 | Alarm management: Centralized alarm display and review with visual/ audible alert prioritization
23 | Bedside parameter view: ECG,SpO2, NIBP,Temperature
24 | System Backup: Regular automated data backup and emergency recovery options
25 | User Access Control: Multi-level user access rights with login/password control
26 | Remote Bedside Control: Change monitor settings from CMS station
27 | Recording: Support for laser/ thermal printers and digital export (PDF/CSV/XML formats)
28 | Waveform Display: Up to 8 waveforms per patient in real time, customizable per user preference
29 | Hardware requirements: Supplied server workstation, mouse, keyboard and networking components
30 | Language support: Multilingual user interface (Including Arabic and English)
31 | Data export: network-based export of waveforms, ECGs, events and reports
32 | Event review: Automatic event recording (Arrhythmias , alarms, NIBP measurements)
33 | ECG management: Automatic arrhythmia analysis, ST segment trend, and 12-lead ECG storage and review
34 | Trend Data: Graphic and tabular trends up to at least 200 hours, trend zoom and review by parameter
35 | Must be supplied with 1 TB HDD or more
36 | Full Disclosure: Up to 10 days of full disclosure waveform storage per patient for the CMS and at least 48 hours of full waveform data per patient for the monitors
37 | Full DICOM and HL7/eGateway compatibility , Integrated with HIS
38 | Specify Dimensions and weight of the unit
39 | Quality and Patient safety
1 | Unit must be FDA / CE marked .
2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
3 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1.
40 | Accessories
Supplied with all standard accessories.
41 | Consumables
Supplied with all standard consumables.
42 | Available Additional Option
No. | OPTIONS DESCRIPTION
43 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .