## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Automated Coagulation Analyzer | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
Technical offer must be clear in terms of standards , which will be offered with the Automated Coagulation Analyzer and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
1 | Test Parameters
1.1 | Routine test (Pothrombin Time, Activated partial thromboblastin Time, Fibrinogen & D-Dimer)
1.2 | Special Coagulation test (Factor assay, Thrombophilia profile, Platelet Function test)
2 | Sample and reagent specifications:
2.1 | Test Volume: Capable of processing 100-200 tests per hour
2.2 | Sample Capacity: Ability to process up to 100-150 samples per batch
2.3 | Sample Volume: Low sample consumption—ideally, 100-200 µL per test
2.4 | Accommodates different tubes type
2.5 | Reagent Shelf Life: Long shelf life for reagents, and easy reagent preparation
3 | Detection Method (Multiple Detection Methods)
3.1 | Mechanical Detection: for routine tests (PT, APTT, Fibrinogen).
3.2 | Automated Preanalytical sample check (HIL system)
3.3 | Wide analytical measurement range.
4 | Calibration and Quality Control
4.1 | Automatic calibration (less calibration frequency, pre-calibrated reagents)
4.2 | Internal QC: long QC material stability
4.3 | External QC Support: Ability to integrate with external QC programs like CAP (College of American Pathologists) (Large peer group)
4.4 | Result Accuracy: Results should meet industry standards (e.g., CLSI guidelines) for accuracy and reproducibility.
4.5 | Internal Quality Control Peer review
4.6 | Comprehensive validation protocol
5 | Connectivity
5.1 | Data Management: Ability to store test results, patient information, and analytical data for prolong period of time (5 years).
5.2 | If middleware Software is used, it needs to have multiple access (can be displayed in different PCs). | 1
5.3 | Interface Requirements for HIS Integration (Track care-TCLE) Protocol Support The machine must support ASTM or HL7 protocols to enable seamless integration with the Laboratory Information System (LIS)/Trak Care Lab Enterprise. Software Components Any required software components, middleware, or workstation setups necessary for integration must be provided by the supplier. Skilled Resources The supplier should have trained and skilled personnel with expertise in LIS/HIS integration to support implementation and troubleshooting. Documentation The supplier is required to provide detailed LIS integration specification documents covering all technical aspects of the interface.
6 | Automation
6.1 | Automated Sample Dilution when applicable.
6.2 | Batch and Continuous Testing: The analyzer should support both batch testing (for routine tests) and continuous testing (for STAT or urgent tests).
6.3 | Ability of the analyzer to be integrated in the total lab automation.
7 | Maintenance and Support
7.1 | Self-Diagnostics: The system should have self-check features to detect malfunctions and ensure the system is operating properly.
7.2 | User-Friendly Maintenance Protocols: Maintenance tasks should be straightforward and well-documented for staff.
7.3 | Technical Support: Availability of prompt on-site or remote technical support to minimize downtime.
8 | Size and Design
8.1 | Compactness
8.2 | Ease to operate
8.3 | Safety Features
8.4 | Software that support all accreditation tools (reagent comparison, analyzer comparison, QC lot verification,..etc.)
9 | Specify Dimensions and weight of the unit
10 | Unit should be provided with all necessary accessories and consumables.
11 | Must have universal power adapter 220-240v ,50/60HZ
12 | supplier must provide the information on energy efficiency options in the equipment, sustainability methods adopted, any carbon footprint technologies utilized, and certifications
13 | Quality and Patient safety
13.1 | Unit must be FDA / CE marked .
13.2 | Number of customers used in CAP accredited labs in UAE / ME Region
13.3 | CAP external quality control program enrollment
13.4 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
13.5 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1.
14 | Sample Consumption
1500 sample per month(5000 test )
15 | Consumables
Supplied with all standard consumables.
16 | Available Additional Option
No. | OPTIONS DESCRIPTION
17 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End user training for lab staff and later as needed .
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS,CAP
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warranty : Comprehensive minimum 5 Years warranty .