## Ultrasound -Neuro
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| PORTABLE VASCULAR ULTRASOUND FOR OPD-Neurology | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Ultrasound machine and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | This premium ultrasound machine delivers top-of-the-line performance, featuring advanced 3D imaging and exceptional image clarity, making it the ideal choice for neurology departments. The system is equipped with high-resolution transducers and cutting-edge imaging software that enable precise visualization of neurological structures, ensuring accurate diagnostics and treatment planning.
3 | This ultrasound machine is a top-tier system designed to support all neurology applications, offering unparalleled image quality and performance. It includes a wide range of features specifically tailored to the unique needs of neurological diagnostics, ensuring precise and reliable results for even the most complex cases. Includes 1)High-Resolution Imaging,2)3D Imaging Technology3 )Doppler Ultrasound 4)Real-Time Imaging and 5)Multimodal Imaging Software
4 | Should have two to four Active Probe slots.
5 | Transducer Types: • Liner high frequency probe • micro convex • Phased Array
6 | Volume Navigation
7 | Must include Operating Modes • B-Mode •Colour Doppler: Anatomical & Colour M-Mode • Colour Flow Mode (CFM) • Extended Field of View • Power Doppler Imaging (PDI) • Pulse Wave Doppler • CW Doppler • Tissue Doppler• •Power Doppler •Spectral Doppler •3D• Colour Doppler
8 | Unit should be fully supported to PACS system-Capability to store and archive images and video clips in various formats (e.g.DICOM, JPEG)
9 | •Monitor Screen minimum 24′′ High-Resolution LCD
10 | specify Min/Max Hight adjustment
11 | Horizontal/Vertical viewing angle of monitor
12 | Image Storage: •Storage Formats:DICOM – compressed/uncompressed, single/multiframe with/ without RawData •Export JPEG, WMV (MPEG 4), and AVI formats • support Storage Devices:USB Memory Stick,CD,DVD, Hard Drive storage
13 | Peripheral Options • User Interface Operator Keyboard • Digital B&W thermal printer • DVD video recorder -(as optional item) •Gel warmer
14 | Connectivity & DICOM : • Ethernet network connection• Verify• Print• Store• Modality Worklist • Storage Commitment• Query / Retrieve
15 | Having B/M color and PW/CW
16 | Unit Must have HDMI /DVI/VGA ports
17 | Frame rate/Sampling rate must not be less then 150 frames/sec
18 | Must be supplied with 2 TB HDD
19 | Full DICOM and HL7 compatibility
20 | Specify Dimensions and weight of the unit
21 | Unit should be provided with all necessary accessories and consumables.
22 | Must have universal power adapter 220-240v ,50/60HZ
23 | supplier must provide the information on energy efficiency options in the equipment, sustainability methods adopted, any carbon footprint technologies utilized, and certifications
24 | Quality and Patient safety
25 | Unit must be FDA / CE marked .
26 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
27 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1.
27 | Accessories
Supplied with all standard accessories.
28 | Consumables
Supplied with all standard consumables.
29 | Available Additional Option
No. | OPTIONS DESCRIPTION
30 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .

## Ultrasound-Ortho
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| PORTABLE ULTRASOUND FOR OPD-Orthopedic | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Ultrasound machine and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | This premium ultrasound machine offers top-of-the-line performance, including advanced 3D imaging and exceptional image clarity, making it ideal for use in orthopedic departments. The system is equipped with high-resolution transducers and sophisticated imaging software that enable precise musculoskeletal assessments, joint evaluations, and soft tissue diagnostics.
3 | Ultrasound machine is a top-tier system designed to support all orthopedic applications, offering unparalleled image quality and performance. It includes a wide range of features specifically tailored to the needs of orthopedic assessments, such as high-resolution imaging for joint, tendon, ligament, and soft tissue evaluation. The system excels in the visualization of musculoskeletal conditions, from routine diagnostic scans to complex injury assessments.
4 | Should have two to four Active Probe slots.
5 | Transducer Types: • Liner high frequency probe • Hocky stick probe • convex probe
6 | Volume Navigation
7 | Must include Operating Modes • B-Mode •Colour Doppler: Anatomical & Colour M-Mode • Colour Flow Mode (CFM) • Extended Field of View • Power Doppler Imaging (PDI) • Pulse Wave Doppler • CW Doppler • Tissue Doppler• •Power Doppler •Spectral Doppler • Colour Doppler
8 | Unit should be fully supported to PACS system-Capability to store and archive images and video clips in various formats (e.g.DICOM, JPEG)
9 | •Monitor Screen minimum 24′′ High-Resolution LCD
10 | specify Min/Max Hight adjustment
11 | Horizontal/Vertical viewing angle of monitor
12 | Image Storage: •Storage Formats:DICOM – compressed/uncompressed, single/multiframe with/ without RawData •Export JPEG, WMV (MPEG 4), and AVI formats • support Storage Devices:USB Memory Stick,CD,DVD, Hard Drive storage
13 | Peripheral Options • User Interface Operator Keyboard • Digital B&W thermal printer • DVD video recorder -(as optional item) •Gel warmer
14 | Connectivity & DICOM : • Ethernet network connection• Verify• Print• Store• Modality Worklist • Storage Commitment• Query / Retrieve
15 | Having B/M color and PW/CW
16 | Unit Must have HDMI /DVI/VGA ports
17 | Frame rate/Sampling rate must not be less then 150 frames/sec
18 | Must be supplied with 2 TB HDD
19 | Full DICOM and HL7 compatibility
20 | Specify Dimensions and weight of the unit
21 | Unit should be provided with all necessary accessories and consumables.
22 | Must have universal power adapter 220-240v ,50/60HZ
23 | supplier must provide the information on energy efficiency options in the equipment, sustainability methods adopted, any carbon footprint technologies utilized, and certifications
24 | Quality and Patient safety
25 | Unit must be FDA / CE marked .
26 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
27 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1.
27 | Accessories
Supplied with all standard accessories.
Sample Consumption
100 microlite or lesss per month
28 | Consumables
Supplied with all standard consumables.
29 | Available Additional Option
No. | OPTIONS DESCRIPTION
30 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .