## Dialysis Machine & RO System
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Hemodialysis Machine | 7 | Comply | Supporting Documents Ref. | by Hospital
2 | Portable RO Unit | 1
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
Item 1: | Hemodialysis Machine (FOC against consumables)
1 | General Data: |  | Yes
1.1 | Please specify make, model, and origin of the machine.
1.2 | Please specify the dimensions of the unit.
1.3 | Please specify the weight of the unit.
1.4 | Power supply: 220 to 240 VAC ±10%, 50 to 60 Hz ±1 Hz (≈10A)
1.5 | Unit should have monitor to display dialysis parameters and graphs. Please specify size of monitor
1.6 | Concentrate supply pressure: 0 to 0.5 bar
1.7 | Water supply:
1.7.1 | Pressure: 1 to 7 bar at minimum 800 mL/min at maximum 3000 mL/min
1.7.2 | Temp.: 5 to 30 °C
1.8 | Drain:
1.8.1 | Minimum drain capacity: 800 mL/min average
1.8.2 | Height: 50 cm maximum
1.8.3 | Temp.: 90 °C maximum
2 | Treatment modalities:
2.1 | HD
2.2 | HDF
2.3 | HF
2.4 | Online HDF
2.5 | Online HF
3 | Hydraulic Circuit:
3.1 | Unit contains Endotoxin retentive filter (ETRF)
3.2 | Dialysis fluid flow rate setting range: 300 to 800 mL/min
3.3 | Dialysis fluid temperature setting range 34.0 to 40.0 °C
3.4 | Dialysis fluid conductivity:
3.4.1 | Dialysate type: Bicarbonate / Acetate
3.4.2 | Bicarbonate dialysis Bicarbonate conductivity setting range: 2.3 to 7.0 mS/cm Accuracy: ±0.1 mS/cm Total conductivity setting range: 12.7 to 15.2 mS/cm Accuracy: ±0.2 mS/cm Acetate dialysis Total conductivity setting range: 12.7 to 15.2 mS/cm Accuracy: ±0.2 mS/cm
3.5 | Transmembrane pressure (TMP) Measurement range: -100 to +500 mmHg Measurement accuracy: ±10 mmHg
3.6 | Unit has Blood leak detector:
3.5.1 | Sensitivity: 0.3 mL Blood/1 L Dialysis fluid (Blood: Hematocrit 32 ± 2%; Dialysis fluid temperature: 37 °C)
3.6 | Ultrafiltration:
3.6.1 | UF rate: 0.00; 0.10 to 4.00 L/h UF accuracy (Balance): ±30 mL/h (At dialysis fluid flow rate 300 to 500 mL/min) ±0.1% of the dialysis fluid flow rate (At dialysis fluid flow rate 501 to 800 mL/min)
3.7 | Dialysis Dose Monitor:
3.7.1 | Measurement principle: Absorptiometry Applicable Treatment mode: HD, On-line HDF Applicable Kt/V range: 0 to 3.0 Kt/V monitoring accuracy: ±0.1 (Kt/V 0 to 1) ±10% (Kt/V 1 to 3) Applicable URR range: 0% to 100% URR monitoring accuracy: ±5%
4 | Extracorporeal circuit:
4.1 | Arterial pressure monitoring: Measurement range: -300 to +500 mmHg Measurement accuracy: ±10 mmHg
4.2 | Venous pressure monitoring: Measurement range: -300 to +500 mmHg Measurement accuracy: ±10 mmHg
4.3 | Dialyser inlet blood pressure monitoring: Measurement range: -300 to +735 mmHg Measurement accuracy: ±10 mmHg
4.4 | Single needle pressure: Measurement range: -200 to +600 mmHg Measurement accuracy: ±10 mmHg
4.5 | Air detector: Sensitivity: 0.02 mL (normal air bubbles) (At Blood flow rate: 250 mL/min) 0.0003 mL (microbubbles: blood/air mixture) (At Blood flow rate: 250 mL/min)
4.6 | Arterial blood pump: Setting range: 40 to 600 mL/min Flow rate accuracy: Set value ±10% (inlet Pressure -150 mmHg P +150 mmHg) Set value -20 to 0% (inlet Pressure -200 mmHg P -150 mmHg)
4.7 | Heparin pump: Setting range: 0.0 to 9.9 mL/h Output rate accuracy: Set value ±10% Syringe type: 30 mL or 20 mL, 20 mL or 10 mL (optional) Bolus volume: 0.0 to 9.9 mL
4.8 | Venous blood pump / Substitution fluid pump: Setting range: 40 to 600 mL/min Flow rate accuracy: Set value ±10% (inlet Pressure -150 mmHg P +150 mmHg) Set value -20 to 0% (inlet Pressure -200 mmHg P -150 mmHg)
4.9 | Blood Pressure Monitor (BPM):
4.9.1 | Pressure display range: 10 to 300 mmHg Pressure display accuracy: Less than ±3 mmHg Measurement range (adults): Systolic blood pressure (SYS): 60 to 250 mmHg Mean arterial pressure (MAP): 45 to 235 mmHg Diastolic blood pressure (DIA): 40 to 200 mmHg Pulse rate: 40 to 200 beats per minute
4.9.2 | For iNIBP: Pressure display range: 0 to 300 mmHg Pressure display accuracy: Less than ±3 mmHg Measurement range: Adult systolic blood pressure (SYS): 40 to 280 mmHg Mean arterial pressure (MAP): 10 to 280 mmHg Diastolic blood pressure (DIA): 10 to 235 mmHg Pulse rate: 30 to 200 beats per minute
4.10 | Unit has Blood Volume Monitor (BVM):
4.10.1 | Applicable blood flow rate range: 40 to 600 mL/min Applicable hematocrit range: 15 to 50% Accuracy: ±2.3 dBV%
4.11 | Vascular access recirculation rate measurement:
4.11.1 | Measurement range: 0~100% Measurement accuracy: ±10 (recirculation ratio %) Measurement times: Maximum 5 times Usable blood flow rate range: 100~600 ml/min Usable hematocrit value range: BVM measurement range 15 to 50%
5 | Cleaning:
5.1 | Disinfection and decalcification: Heat disinfection with 50% Citric acid Chemical disinfection with DIALOX (Peracetic acid)
5.2 | Disinfection and degreasing: Sodium hypochlorite solution (Maximum 10%)
5.3 | Decalcification: 30% Acetic acid
6 | FDA approved / CE Marked
7 | Offered System must comply with Electrical safety IEC 6060-1
8 | Device must be manufactured on quality standards. Medical device Quality management system , ISO-13485 or equivalent.
Item 2: | Portable RO Unit (FOC)
1 | 230 V / 50-60 Hz Supply Voltage
2 | Dialysis Water Pressure: 3 bar
3 | Membrane Retention Rate: >95%
4 | Flow in L/min of the unit should be at least 1.5 L/min (90LPH) of product water with stable pressure
5 | Unit must have wheels and be easily movable
6 | Unit must have maintenance alarms that indicate if there are any faults or errors within the machine
7 | Please specify the type of filters used by the unit and how often it must be changed
8 | Unit must not be bulky and should be able to fit in small spaces
9 | FDA approved / CE Marked
10 | Offered System must comply with Electrical safety IEC 6060-1
11 | Device must be manufactured on quality standards. Medical device Quality management system , ISO-13485 or equivalent.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 5 Years warrenty .