## Fluroscopy Unit
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Digital Fluoroscopy unit | 1 | Comply Yes/No (If Yes Specify) | Supporting Documents Ref. | by Hospital
Specifications
The start -of -the-Art Digital Fluoroscopy system with maximum 80-90 kW ,900 mA,17x17 FDP detectors, To minimize X-ray dose and improves Examination workflow , Single Auto Stitching ,Quick preview after exposure ,floor mounted ,Tilting table and Auto positioning features. Ceiling Suspension Monitor for flexible, comfortable, and safe positioning and Examination light .
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | All pre-installation requirements should be submitted along with the quotation including suggested room dimensions etc.
4 | Integration with PACS , GE Dose Watch & RIS :Unit must be able to integrate with hospital PACS system & GE Dose Watch .All software's and hardware's must be provided by vendor to do the integration .
4.1 | DICOM Image Storage Service Class as defined below: The DICOM Storage Commitment Service Class For retrieving prior images from the PACS, the DICOM Query/Retrieve.
5 | X-ray Tube & Monitor Suspension:
5.1 | Dual focus X-ray tube( 0.6 / 1.3 mm) with high anode heat capacity and cooling rate
5.2 | Output power of 80-90 KW
5.3 | Fluoroscopy output with PCF(Pulse controlled Fluoroscopy ) 125kV & GCF (Grid controlled fluoroscopy) 110 kV or better
5.4 | Floor suspension with motorized height adjustment
5.5 | Motorized Tube column Angulation , support Oblique exposures at both ends of tabletop possible
5.6 | Motorized Tube column longitudinal Movements slow and fast allow Patient coverage without patient movement
5.7 | Control handle with automatic collimation with timer
5.8 | Collimation LED light switches automatically on when needed
5.9 | SID positioning with laser beam & tape line
5.1 | Digital SID min 180 cm
5.11 | Protection Device like Broken filament, Abnormal rotor, Overload and High voltage arcing
5.12 | The mAs are displayed according to the constant chosen an optional display of the mAs accumulates during the exposure the mAs are also displayed during an exposure when using the automatic exposure controller
5.13 | Error code display: 1.Tube overheating. 2.High voltage, power supply not ready. 3.Exposure request which exceeds the max power the generator exposure request for too many mAs for the exposure. 4.The exposure is impossible due to the absence of the emission characteristics of the tube
5.14 | Automatic Positioning for X-ray Tube & Detectors
5.15 | Reset to last SID position
5.16 | Electromagnetic brakes in all directions & central release feature
5.17 | Sealing mounted Dual medical grade Monitor suspension up to 22" or Two in one monitor up 32''
5.18 | Sealing mounted procedure light and Ceiling mounted Lead shields
6 | Digital Radiography Mode
6.1 | kVp range 40 - 150 or better
6.2 | No. of kVp increments +/- 1kv
6.3 | mA range up to 1000 or better
6.4 | Please specify No. of mA increments
6.5 | Please specify mAs range
6.6 | Please specify No. of mAs increments
6.7 | AEC
6.8 | Exposure time range in (sec)
6.9 | Please specify No. of time increments
6.1 | Anatomical programming
6.11 | Acquisition mode continuous fluoroscopy
6.12 | kVp range up to 40-150 or better
6.13 | mA range 40 - 125 or better
6.14 | ABS control (kVp and mA)
6.15 | Maximum exposure rate (R/min) 10 or better
7 | Acquisition Mode Pulsed Fluoroscopy
7.1 | X-ray Tube (Fluoroscopy) 20 or better
7.2 | Tube Power Rating (kW)
7.3 | Nominal kVp 80 or greater.
7.4 | Anode Heat Storage Capacity (kHU) range:300-800kHU
7.5 | Heat Dissipation Rate (kHU/min)
7.6 | Grid switch mechanism for pulse fluoroscopy 300 or greater.
7.7 | Cooling System Type ?
7.8 | Please specify Focal Spot
7.9 | Please specify Small (mm)
7.1 | Please specify Medium (mm)
7.11 | Please specify Large (mm)
8 | Detector Size 43x43 cm(17x17 inches)
8.1 | Active Imaging Area-Not less than 42x42 cm
8.2 | Specify DQE
8.3 | Specify MTF
8.4 | Resolution 2,500 x 2,500 ( Best preferred)
8.5 | Pixel Size (mm) <148 micron (lower preferable)
8.6 | Scintillator(Csi)
8.7 | Acquisition depth (pixel depth) 16 bits
8.8 | Maximum acquisition speed (im/s) up to 30
8.9 | Maximum diagonal size (cm) 43x43
9 | Fluoroscopy Digital Bucky Table:
9.1 | Height adjustment range min 60- 140 cm
9.2 | Adjustable table height- 90° / + 90° tilt range
9.3 | Specify Tabletop travel range : 1)longitudinal 2)Transverse
9.4 | Max. patient load 250 to 300 Kgs
9.5 | Balanced at counter weight and table movement: 4 ways with breaks
9.6 | Removable footrest
9.7 | Motorized Compression
9.8 | Maximum compression force (KG) Up to 15 or better
9.9 | Load bearing capacity up to 250Kg or more
9.1 | Collision prevention sensors for downward movement
9.11 | Cassette sensing feature
9.12 | Easy to clean surface
9.13 | Foot & hand switch for operation
9.14 | Emergency power off switch
10 | Control & Image Pre-Processing Console
10.11 | Floor & Ceiling mounted type
10.12 | Compact control console with display screen
10.13 | Organ programming features
10.14 | Manual & automatic parameter settings
10.15 | Automatic exposure control (AEC) 3 chambers
10.16 | Post exposure data & Under exposure indication
10.17 | Tube protection against over load
10.18 | Anatomically specific pre-processing
10.19 | Auto ranging feature
10.2 | Integrated DAP Monitor with shortest warm-up time
10.21 | Exposure Index feature
10.22 | Archival Storage: Local Hard Drive capacity > 1 Tera Byte, RW/+RW/-DL/Read, CD-R, DVD,+RW support for Image storage, ≥23’’ HR flat panel 6 MP High resolution medical grade monitor with status and error messages on in-built/1 external PC monitor
10.23 | Fully customizable patient work list
10.24 | Software as required with no additional cost
10.25 | Upgradeable to higher version
10.26 | The no. of mouse clicks to be performed to get an image should be as minimum as possible, but should not exceed 3 mouse clicks
10.27 | Touch screen for fast and intuitive operation in a dynamic X- ray environment
10.28 | Full screen viewing mode for fast decisions
10.29 | 1, 2 , 3 to 4 image display for image comparison
10.3 | The patient name should disappear from the patient display automatically after the first exposure for privacy reasons
10.31 | Auto/Manual Pre-filter selection
10.32 | A separate workstation for image positioning and patient demographic data is required and the workstation should be able to send and receive DICOM (Digital Imaging and Communications in Medicine) standards
11 | Software:
11.1 | Repeat /Reject Analysis software
11.2 | Statistical tool software
11.3 | Protocol analysis
12 | Unit must be antimicrobial Protection-Easy to clean surface
13 | Specify Weight and Dimensions of the unit
14 | Power consumption & Heat dissipation in BTU/h
15 | Accessory:
15.1 | All start-up accessories as per end user to be included
15.2 | Half Mobile Lead apron on wheels x 1 ,Patient foot rest, shoulder supports, patient hand grips, compression band, Armpit support.
15.3 | Half lead apron x 1
15.4 | Whole Spine Imaging accessories
15.5 | Mobile Detector holder ,bed holder
16 | Trade-in (If relevant)
16.1 | Supplier should offer a Trade-in price for the following existing system:
16.2 | DESCRIPTION MANUFACTURER MODEL YEAR OF PRUCHASE
16.3 | FLUOROSCOPY SYSTEM SIEMENS, GERMANY LUMINOS DRF 2011
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
17.4 | Quality control ,Quality assurance and acceptance test should be done by the company
17.5 | Safety assessment should be done by the company
18 | Accessories
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Available Additional Option
| OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
21 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 10 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at least 5 years comprehensive from date of installation and equipment acceptance.
12 | Vendor Must include Site preparation (pre-Installation) work for proposed model

## Civil & MEP Works
Civil and MEP works for General Purpose X-ray Machine
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Civil and MEP works for Fluoroscopy unit | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
All Pre-installation requirements should be submitted, including suggested room Design and dimensions
1 | All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required.
2 | Scope of Proposed Design and As built Drawings ( after completion of Project)Should be Included
3 | Interior Designing and Decorative works to be Proposed and should Include in scope of Works
4 | Demolition and Mobilization works should be included if required
5 | Mobilization(Wall opening and closing ) for X-ray machine delivery should be included
6 | The system offered shall be designed to operate normally under the conditions of UAE, Environmental factors, FNAR and MOH policy the conditions include Power Supply,Climate,Temperature, Humidity,Dust,etc
7 | Detailed Lead test reports for existing lead and compliance certification to be provided at each step to ensure the radiation leakage not only meets but also exceeds the manufacturers requirements.
8 | Provide overlap of the material at each panel joint for the maximum level of performance, attenuation integrity and electrical conductivity.
9 | To install appropriate internal fit out or lining, lighting and electrical meet international standards.
10 | supply and installation of Warning light and signs
11 | Partition works and Ceiling works should be Similar to UHS standards. Proposed Specification need to be provided
12 | Aesthetic Ceiling picture can be done in X-ray room as per client suggestion
13 | Acoustic ceiling tiles.
14 | The Sound insulation can be incorporated into the framework of the walls during assembly and to ensure standard acoustic internal lining.
15 | Wall treatment
16 | To provide clearest observation window with observation glass
17 | The observation Glass shall be anti-reflective coating
18 | Flooring- Medical Grade (similar or better than existing flooring at UHS)
19 | Removing of existing flooring & levelling with self-levelling screed for complete area.
20 | Power Supply Capacity and Specifications
21 | Proposal for Supply and Installation of Electrical works. socket and switches should be via brand and of same type used in UHS premises.(All wires will be concealed and in conduct pipes and the wiring for electrical and network cable will be commensurate with the existing ).
22 | The lighting should be all LED and should support energy conservation
23 | HVAC- Study the current layout and Propose the modifications required
24 | Fire Alarm and Fire Fighting modification works to be propose
25 | Supply and installation Bedhead units for proposed locations including Medical Gas Points and Nurse call system
26 | Existing ceiling mounted procedure light and ceiling mounted lead shield can be used.
27 | All MEP work including medical gas system should be according to UHS stander which included future line if needed and current system assessments to avoid any struggling with operation after the installation
28 | lead lining should be according to the FANR requirements As per the safety assessment ,Tube outcome ,Generator power and machine orientation.
29 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
30 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Equipment. Two Copies each and one Soft Copy for each.
3 | It is mandatory for the Contractor to submit End-of-Sale, End-Of-Support, and End-Of-Life for each individual hardware component - Proof documents from the manufacturer to be attached with the proposal.
4 | Proposal Should be submitted after having site survey and comparison with the current As built
5 | All The works Should Comply with Standards of Civil defense, SEWA and MOH
6 | Contractor should discuss the final technical proposal with the technical team before submission. Proposal should include ongoing warranty, support.