## NICU ventilator
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Portable Neonatal/Pediatric Ventilator for NICU application | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Neonatal/Pediatric Ventilator and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | Please specify make, model, and origin of the machine.
3 | Ventilation modes: Volume targeted modes, Pressure controlled modes, Intelligent ventilation, invasive and Non-invasive modes
4 | Standard functions: CPR ventilation, event log, flow trigger, manual breath, O2 enrichment, SpO2 monitoring, USB port, ethernet port ,Serial port for HIS
5 | Monitoring parameters: pressure, flow, volume, oxygen, time, lung mechanics, CO2, SpO2.
6 | Display Waveforms
7 | High, medium, low priority Alarms
8 | Specify Battery backup
9 | Ability to perform daily self test by end-user
10 | Unit must be compatible with Nitric oxide delivery system(SLE3600 INOSYS)
10 | General Specification:
10.1 | Patient Group: Neonates, Paediatrics.
10.2 | 15-inch or more Detachable Colour Touchscreen.
10.3 | Time-controlled, flow-controlled, pressure-controlled, volume-controlled
10.4 | Integrated High-Frequency Oscillation (HFO)
10.5 | Integrated Heated Patient Gas Humidification System or external Humidification system.
10.6 | Integrated non-invasive ventilation and high-flow oxygen therapy.
10.7 | Dead Space free Flow Sensor, to be used during Invasive and NIV.
10.8 | Integrated fully automated Oxygen saturation controller for real-time oxygen saturation adjustment.
10.9 | With Mobile stand (Trolley)
11 | CO2 Measurement
12 | Ventilator Should be modular and capable to add any upcoming new monitoring parameters. (Software or Hardware)
13 | Trigger – Flow, Pressure, External.
14 | Unit should have One step calibration procedure after patient breathing circuit setup.
15 | Ventilation Modes:
15.1 | Pressure-Controlled (PC): PC-CMV, nPC-CMV, PC-SIMV, nPC_SIMV, PC-SIMV+,nPC-SIMV+, PC-ACV, nPC-ACV, PC-ACV+, nPC-ACV+, CPAP, SnCPAP, PAV, nPAV, HFO, nHFO, SnHFO, High-Flow, manual.
15.2 | Volume-Controlled (VC): VC-CMV, VC-SIMV, VC-ACV, VC-ACV+
15.3 | Volume limitation, leakage compensation, PSV, PAV.
15.4 | Manoeuvres: Inspiration hold/manual, preoxygenation, medication nebulization.
16 | Ventilation Parameters:
16.1 | Inspiration Pressure (Pinsp): 5 - 95 mbar
16.2 | Inspiration time Ti: 0-30 s
16.3 | PEEP (End Expiratory Pressure): 0 - 40 mbar
16.4 | Tidal Volume: 1 - 2000ml
16.5 | Respiratory Rate (f): 1 - 300 breaths/min
16.6 | FiO2 (Fraction of Inspired Oxygen): 21 - 100%
16.7 | Compliance (C): 0 - 900 ml/mbar
16.8 | Resistance (R): 0 - 500 mbar/l/s
16.9 | Patient Gas Temperature (near patient): 10 - 50°C
17 | High-Frequency Oscillation (HFO) Parameters
17.1 | Frequency: 5 - 25 Hz
17.2 | Mean Airway Pressure (MAP): 0 - 40 mbar
17.3 | Oscillatory Amplitude (Posc): 1 - 180 mbar
17.4 | Inspiration Ratio: 30 to 50%
18 | Monitoring & Alarms
18.1 | Airway Pressure: High/Low
18.2 | Expiratory Minute Volume (MVe): High/Low
18.3 | Expiratory Tidal Volume (VTe): High/Low
18.4 | End Expiratory Pressure (PEEP): High/Low
18.5 | Mean Airway Pressure (MAP): High/Low
18.6 | Respiration Rate: High
18.7 | Apnea, Leakage, Disconnection, Occlusion Alerts
18.8 | CO2 Monitoring (EtCO2) keep as optional
18.9 | FiO2: High/Low
18.10 | Base FiO2: High
19 | Specify Dimensions and weight of the unit
20 | Unit should be provided with all necessary accessories and consumables.
21 | Must have universal power adapter 220-240v ,50/60HZ
22 | Emergency Power: Min. 60 minutes on rechargeable batteries
23 | Gas Supply: Air: 2.8 - 6.0 bar and Oxygen (O2): 2.8 - 6.0 bar
24 | Quality and Patient safety
24.1 | Unit must be FDA / CE marked .
24.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
24.3 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
25 | Consumables
25.1 | Should be Supplied with all standard consumables.
26 | Accessories
26.1 | Should have All standard Accessories must be included as per manufacturer .
27 | Available Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training if or operator and Physians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation

## NICU CPAP ven
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Portable Neonatal/Pediatric CPAP Ventilator for NICU application | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the CPAP Ventilator and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
1.1 | Please specify make, model, and origin of the machine.
2 | Equipment should be suitable for both, neonates and pediatric patients
3 | Specify Battery backup and Serial port for HIS
4 | General Specification:
5 | Non-Invasive Ventilation modes (NIV) for neonates and peadiatric patients
6 | Device should flow variable and pressure constant
7 | Built in battery backup (At least 4 Hrs.)
8 | Color TFT LCD touch screen display
9 | O2 therapy with High Flow Nasal Cannula
10 | Should be included with the following modes:
11 | Nasal intermittent positive pressure ventilation (NCPAP)
12 | NIPPV (nasal IPPV)
13 | SNIPPV (Synchronized NIPPV)
14 | High-flow nasal cannula (HFNC)
15 | NIV Trigger system
16 | Leak Compensation upto 25%
17 | Flow- 0.5-20 L/min
18 | Ti- 0.1-20 s
19 | Respiratory Rate- 1-120 bpm
20 | T apnea- Off, 10-30s
21 | Cpap Pressure- 1-13 cmH2O
22 | PEEP- 1-13 cmH2O
23 | Pinsp- 3-20 cmH2O
24 | Manual Ventilation time – 20s
25 | Manual Breathing option available- 3-20 cmH2O, 3-25 L/min
26 | Flush O2- 23-100%, Time- 30-120 s
27 | Specify Dimensions and weight of the unit
28 | Unit should be provided with all necessary accessories and consumables.
29 | Must have universal power adapter 220-240v ,50/60HZ
30 | Emergency Power: Min. 60 minutes (without humidifier) on rechargeable batteries
31 | Gas Supply: Air: 2.0 - 6.0 bar and Oxygen (O2): 2.0 - 6.0 bar
21 | Quality and Patient safety
21.1 | Unit must be FDA / CE marked .
21.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
21.3 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
22 | Consumables
22.1 | Should be Supplied with all standard consumables.
23 | Accessories
23.1 | Should have All standard Accessories must be included as per manufacturer .
24 | Available Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training if or operator and Physians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation

## Transport Incubator with Ventil
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Transport Incubator with Ventilator for NICU | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
| Please specify make, model, and origin of the machine. |  | Yes
1 | Incubator and ventilator:
1.1 | Unit must be integrated with ventilator
1.2 | All incubator controls, ventilator controls, tank gauges, gas-related controls must be front-mounted for quick access and easy to see
1.3 | oxygen cylinders with holders must be housed in the unit's frame beneath the incubator for easy transport
1.4 | Gas tank regulators, yoke fittings, and flowmeters must be built-in which eliminates the need for external tubing
1.5 | Unit must be able to run on battery power for at least 4 hours
1.6 | Unit must have double walls for safety
1.7 | Unit comes with skin temperature probe
1.8 | Unit has O2 inlet
1.9 | Unit has access doors on each side
1.10 | Unit has 2 port holes on each side
1.11 | Unit has an O2 concentration range between 21% to 100 %
1.12 | Unit has humidity capacity between 50%and 70%
1.13 | Unit noise level must be less than 60dBA
1.14 | Unit has a temperature set range between 22°C to 38°C
1.15 | Unit has an air filter which removes >99% of airborne particles greater than 0.5 micron diameter
1.16 | Unit has battery indicators to indicate battery life percentage
1.17 | Unit has indicators to inform user if the unit is running on AC power or DC power
1.18 | Unit has infant temperature display
1.19 | Unit has indicators to indicate the heater power levels
1.20 | Unit displays incubator air temperature
1.21 | Unit displays set temperature
1.22 | Unit must have alarm indicators for high temp, power fail, sensor fault, air flow, low DC
2 | Stretcher:
2.1 | Unit must be compatible with Ferno floor mounted bracket locking system in our ambulance
2.2 | Unit must have a load capacity of at least 160 kgs along
2.3 | Unit must be able to adjust its height between approximately 20 cm to 100 cm
2.4 | Unit must have a length of at least 160 cm
2.5 | Unit must have a width of at least 55 cm
2.6 | Unit must be able to be handled easily without the need for too many helpers
2.7 | Unit must be able to easily be tilted and height adjusted
3 | Specify Dimensions and weight of the unit
4 | Unit should be provided with all necessary accessories and consumables.
5 | Must have universal power adapter 220-240v ,50/60HZ
6 | Emergency Power: Min. 60 minutes on rechargeable batteries
7 | Gas Supply: Oxygen (O2): 2 - 6 bar
8 | Quality and Patient safety
8.1 | FDA approved / CE Marked
8.2 | Offered System must comply with Electrical safety IEC 6060-1
8.3 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
9 | Accessories
Supplied with all standard accessories.
10 | Consumables
Supplied with all standard consumables.
11 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training if or operator and Physians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation