## Wide Field Fundus Camera
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Wide Field Fundus Camera for ophthalmology | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
| Please specify make, model, and origin of the machine. |  | Yes
1.1 | Must provide at least 200 degrees field of view of the retina in one or two shots
1.2 | Images must have a resolution of at least 20 μm
1.3 | Device must include the following image modalities:
1.3.1 | red / green color
1.3.2 | Color Composite
1.3.3 | Red-free
1.3.4 | Choroidal
1.3.5 | Autofluorescence
1.3.6 | Fluorescein Angiography
1.4 | Minimum pupil diameter of at least 2 mm
1.5 | Device must have a fast exposure time to allow for quick exams
1.6 | Please specify working distance
2 | Device must be wheelchair accessible with adjustable bench
3 | Unit must be supplied with color printer
4 | Connectivity:1)Data Management: Integration with hospital information systems (HIS) 2)Data Export: Ability to export results in PDF 3)WirelessConnectivity: Wi-Fi, Bluetooth, or USB for transferring data to external systems
5 | Environmental temperature
6 | Specify dimension of the unit
7 | Weight of the unit
8 | Quality and patient safety
8.1 | FDA approved / CE Marked
8.2 | Offered System must comply with Electrical safety IEC 6060-1
8.3 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
9 | Accessories
Supplied with all standard accessories.
10 | Consumables
Supplied with all standard consumables.
11 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
12 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .