## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Higher End ICU Beds (36'' wide frame) | 3 | Comply | Supporting Documents Ref. | by Hospital
2 | Higher End ICU Beds (40'' wide frame) | 1
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Bed should be FDA/CE Approved
3 | Bed should be made at good quality standards in terms of Efficiency , Quality , safety and Connectivity .
4 | Electrical Bed with fully integrated Air mattress designed for intensive Care Areas, with Battery Backup
5 | Connectivity ready with Nurse Call system
6 | One button function for Bed up/down, backrest up/down, knee up/down, Trendelenburg/revs.Trend
7 | In Bed Scale weighing in Kg/lbs.
8 | 30°/45° Head Angle Alarm
9 | Retractable Foot Control
10 | Full Chair Position with possibility of chair Egress (sit to stand option)
11 | StayInPlace Patient Migration Management System
12 | Alternating Low Pressure (ALP) & Continous Low Pressure (CLP)
13 | Pulmonary Surface (CLRT & CPT)
14 | Full Chair Position with Sit-to-Stand and StayInPlace Technology
15 | Continuous Lateral Rotation Therapy (CLRT)
16 | Chest Physiotherapy (CPT)
17 | Nurse Call, Univ TV & Lighting
18 | Both sides Siderail Nurse Call
19 | Three Mode Bed Exit Alarm to Nurse Call
20 | 5th Wheel Steering
21 | IntelliDrive Powered Transport Steering
22 | Transport Shelf
23 | Permanent IV Pole
24 | Have percussion and vibration therapies in the mattress to help clearing secretions from the lungs
25 | Lateral patient Turn function
26 | Central brake system and steering mechanism located on either side of the bed with clearly marked pedals. With Brake not set Alarm
27 | Should have Auto Contour Feature which allows the bed to prevent the patient from sliding down in the bed.
28 | Caregiver Siderail Controls on both sides of the bed
29 | All Frame Functions Electrical
30 | Split Side Rails, with easy Folding down and One Hand Release Siderails
31 | Radiolucent Sleep Deck, with X-Ray sleeve under patient’s chest, and C-Arm compatibility
32 | Safe Working Load minimum 250 Kgs
33 | Color Touch Screen Interface with digital angle indicator
34 | Roller Bumper System
35 | Angle Indicator On both sides
36 | Night Light
37 | 4 Sockets for I.V pole
38 | Bed not in Lowest position indicator
39 | Advanced Microclimate Technology in the mattress
40 | Head & Foot Board
41 | Drainage Bag Holders
42 | Line Manager
43 | Transport Handles
44 | 2 Vertical Oxygen Tank Holders (both Grey Color for Aluminum Cylinders size: E / 24 Cubic Feet)
45 | Hands Free Emergency CPR & Trendelenburg
46 | Obstacle Detect System at Both Sides of Frame
47 | Caregiver Pendant & Caregiver Foot Controls
48 | All positions must be lockable with secure key
A. | Articulation specification approx.
B. | Head 0°-67°
C. | Thigh 0°-30°
D. | Foot 0°-70°
E. | Trendelenburg 13°
F. | Reverse Trendelenburg -13°
49 | FlexAfoot Retractable Foot Control
50 | Wireless Connectivity
51 | Integrated patient mattress is required ( Powered )
52 | Which address immobility associated complications. (b) Help Protect Skin integrity ( c) Promote patient Comfort .
53 | The unit must comply with all applicable ISO/CE/MDD 93/42 or equivalent.
54 | The unit must comply with quality standard including ISO 13485 or equivalent .
55 | The Unit must be hospital grade , that meets EN 60601-1 performance and safety requirements
56 | Accessories
Supplied with all standard accessories.
57 | Consumables
Supplied with all standard consumables.
58 | Available Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .