## PRP anaylzer
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| POCT CRP(C-reactive protein) ANALYZER | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Point of care CRP (C-reactive protein) testing is requested by the pediatric team. As CRP testing is crucial for promptly diagnosing and treating infections or inflammatory conditions. Introducing CRP tests as a POCT can significantly expedite treatment by providing quicker results compared sending the sample to the laboratory. Importantly, CRP POCT often requires only a simple finger prick rather than blood draw, thereby reducing the stress and discomfort associated with sample collection for young patients. This approach not only speeds up the diagnostic process but also enhances the overall patient experience by minimizing procedural discomfort.
2 | Technical offer must be clear in terms of standards , which will be offered with the fully CRP analyzer. The analyzer should measure-Creative protein to check for inflammation in the patient body .Please specify the details clearly and all the requested specs information must be well supported with reference documents.
3 | Size: Small/ handheld device.
4 | Usage: Easy to use
5 | Software: Ability to view the device from a software or by the connection of a middle ware (AQURE).
6 | Maintenance: Auto-maintenance.
7 | Operator Lock-out (Multiple users)
8 | Calibration due lockout
9 | Automatic Quality control lockout
10 | Certification: FDA or CE certified
11 | Test duration: < 10 minutes
12 | Quality control: Manufacturer Quality control availability
13 | External Quality control: Listed in CAP PT
14 | Sample type: Finger prick (Availability of capillary tube with the kit)
15 | Sample volume: 100 microlite or less
16 | Automatic barcode identification of samples
17 | User friendly ,Easy loading and unloading of samples
18 | Easy maintenance and trouble shooting
19 | LIS integration option if available
20 | Unit must be antimicrobial protection
21 | Specify dimension & weight of the unit
22 | Unit should be provided with all necessary accessories and consumables
23 | Unit must have universal power adapter 220-240v, 50/50HZ
24 | Specify the power consumption for the offered units and heat dissipation(BTU/hr)
25 | Availability for demo unit for demonstration
26 | Quality requirements
a | Unit must be FDA/CE certified
b | Number of customers used in CAP accredited labs in UAE / ME Region
c | CAP external quality control program enrollment
d | Unit must comply ISO 13485 quality management system for medical device or equivalent
27 | Sample Consumption
100 microlite or lesss per month
28 | Accessories
Supplied with all standard accessories.
29 | Consumables
Supplied with all standard consumables.
30 | Available Additional Option
OPTIONS DESCRIPTION
31 | Important Notes & Obligatory Conditions
A | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End user training for lab staff and later as needed .
C | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS,CAP
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warranty : Comprehensive minimum 5 Years warranty .