## Mobile X-ray
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Digital Mobile X-Ray unit | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
Mobile DR X-ray Machine with sliding column, Maximum 30-40 kW ,500 mA,17x17,10x12 Wireless detectors with high image resolution up to 4 Lp/mm , Nano model with fast image transmission capacity
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | Integration with PACS , GE Dose Watch & RIS :Unit must be able to integrate with hospital PACS system & GE Dose Watch .All softwares and hardwares must be provided by vendor to do the intergation .
3.1 | A fully digital radiography system capable of detector exposure in vertical, horizontal and oblique positions to perform general radiography. The unit should be completely integrated (integrated Generator and Image Acquisition) and comprise the following along with auto quality control features incorporated.
4 | X-ray Tube and Generator
4.1 | X-ray tube with high anode heat capacity & cooling rate (Min250W)
4.2 | Specify Heat storage capacity of anode Minimum(Heat Storage: 30000J)
4.3 | Power rating, kW @ 100 kVp
4.4 | X-ray power ≥ 80 kW
4.5 | Tube voltage:40-140 Kv with 1.0kv increment
4.6 | mAs range:0.1 – 500 mAs, 42selections
4.7 | Exposure times :0.001-1.25 s
4.8 | Focal spot (S/L):0.3/1.0
4.9 | PACS & DICOM interface for Image store, Query/Retrieve, Print, Work List Generation, MPPS, and archive
4.1 | mA Range: 30~130mA, variable
4.11 | Console & Display: kV & mA & mAs/sec digital display Density (for human) digital display Prep. & Exposure LED Collimator lamp Switch APR Store Switch
4.12 | Type: Fixed anode Focal spot size: 1.8mm Inherent Filtration: 0.5mm Al Target angel: 15° Collimator: Manual: PCM100 Lamp: 24V150watt, halogen lamp, built-in: 30sec. Timer (DC: 30sec) X-ray Field Size: 35cm * 35cm @ 65cm SID - Charge input: 100V – 240 VAC, 50/60Hz
5 | Tube Movement
5.1 | Horizontal, cm (in)
5.2 | Vertical, cm (in)
5.3 | Rotation
5.4 | Z-axis, °
5.5 | X-axis, °
5.6 | Aluminium Filter mm
6 | Wireless portable detector
6.1 | Type :CsI(Scintillator Cesium iodide)
6.2 | Housing:Carbon fiber
6.3 | Detector sizes:17 x 17 cm; 10 x 12 cm
6.4 | High Image resolution up to 4 Lp/mm
6.5 | Secify weight of each FPD
6.6 | Pixel size: lessthan < 148 µm
6.7 | Matrix Depth 16 bit/Pixel and matrices of 4288×4288
6.8 | FPDs must be IPX7 or equlivent water proof standard
6.9 | Specify number of physical drop allowed for FPDs
6.1 | wireless router
6.11 | The resolution should be minimum of 3.0 lines pair/millimeter. - The pixel size should be 130 um of less. - Detector Quantum Efficiency (D.Q.E) should be more 60% @ Zero Line pairs. - The active matrix size should be 3k x 3k or more.
6.12 | 17.5 MegaPixel sensor in terms of the number of Dedector Eliments
7 | Type of power source
7.1 | AC line
7.2 | Battery type
7.3 | Separate batteries for drive and generator control
7.4 | Energy capacity, mAs @ 80 kV?
7.5 | power of charging -Must have single phase 230V/50-60Hz
7.6 | Charging time, full charge, hr?
7.7 | Can also make exposures using on AC power?
7.8 | Number of images per battery charge
7.9 | Time to display image, sec
7.1 | LineVoltage Compensator
7.11 | Protection level against electric shocks: type B
7.12 | Error code display: - Tube overheating - High voltage, power supply not ready - Exposure request which exceeds the max power the generator exposure request for too many mAs for the exposure - The exposure is impossible due to the absence of the emission characteristics of the tube
8 | Clinical workflow
8.1 | The system must have grate forward visibility and easy maneuvers to support comfortable operation
8.2 | Must have large touch panel display for quickly viewing images
8.3 | specify system start up time
8.4 | Must have Large Image Storage Capacity
8.5 | Shock-resistant body design
8.6 | Fast and easy positioning
8.7 | system access Numeric keypad / PIN code
8.8 | Upgradeable to higher version
8.9 | footprint of the machine (LxBxH)
8.1 | User friendly operator console
8.11 | Hand exposure control
8.12 | SID RANGE, cm (in)
8.13 | Integrated DAP Monitor with shortest warm-up time
8.14 | Focal point distance from floor 55 to 205cm:
9 | Unit must be antimicrobial Protection
10 | Specify Weight and Dimensions of the unit
11 | Power consumption & Heat dissipation in BTU/h
12 | High Imaging Resolution Unit for Medical Mobile Application
13 | Quality control ,quality assurance and accaeptance test should be done by the company
14 | Safety assessment test should be done by the company
15 | Unit must be available on demo basis
16 | Accessories
16.1 | All necessary Accessories, consumables and disposables for starting-up.
16.2 | General Radiology QC tools for performance verification done by Enduser, if required.
16.3 | Lead aprons Quantity ( double face 0.35 mm lead equivalent Adult 2 and Light weight paediatric use 1/2 apron 3 Numbers
16.4 | Thyroid shield protection Quantity 3
16.5 | DICOM Image Storage Service Class as defined below: The DICOM Storage Commitment Service Class
16.6 | For retrieving prior images from the PACS, the DICOM Query/Retrieve (C-FIND and C-MOVE for Patient Root and/or Study Root Model) Service Class should be support.
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
17.4 | User-friendly operation and smooth movement - One-unit package with in-built battery/generator system - Durable with operational implication for different position and parts of the body.
18 | Accessories
18.1 | Supplied with all standard accessories.
19 | Consumables
19.1 | Supplied with all standard consumables.
20 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at laest 3 years comprehensive from date of installation and equipment acceptance.

## C-Arm
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | C-Arm Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
High frequency FDP detector intra operative surgical 2D C-Arm Machine ,DSA(Digtal subtraction angiography) with 29'' Dual ultra high resolution medical grade monitor 6 MP ,exceptional Image Quality and Efficient Maneuverability.
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | A mobile imaging unit used primarily for fluoroscopic imaging during surgical and orthopedic procedures.Consists of a computer used to view, manipulate, store and transfer the images to fuji PACS.
4 | Integration with PACS & RIS :Unit must be able to integrate with hospital PACS system .All softwares and hardwares must be provided by vendor to do the intergration . For retrieving prior images from the PACS, the DICOM Query/Retrieve (C-FIND and C-MOVE for Patient Root and/or Study Root Model) Service Class should be support.
5 | Generator & X-ray Tube
5.1 | Microprocessor controlled high frequency digital generator
5.2 | Continuous, Intermittent & Pulsed fluoroscopy modes of operation
5.3 | Pulsed fluoroscopy ≥ 15 fps
5.4 | Auto & Manual control of fluoroscopy
5.5 | Auto High penetration for thicker patients.
5.6 | Automatic & Manual dose rate control
5.7 | X-ray tube with high anode heat capacity & cooling rate
5.8 | X-ray power ≥ 15kW
5.9 | User friendly operator console
5.1 | Integrated, Detachable, & foot switch exposure control
6 | Mobile C-Arm
6.1 | Easily maneuverable with locking facility
6.2 | wireless router
6.3 | Counter balanced with multi-directional positioning
6.4 | Maximum pivotal & orbital rotation
6.5 | Motorized height & rotational movement with safety control
6.6 | specity C-arm depth
6.7 | Specify Angulation
6.8 | specity lowest lateral postion range
6.9 | Source-Image distance ≥ 90 cm or reasonable
6.1 | Iris Collimator to limit the field to the area of interest which helps reduces radiation exposure to the patient
6.11 | Control of mA and kV parameters
7 | Detector & Image Display
7.1 | High resolution CsI digital panel detector ≥26x26 cm
7.2 | Pixel size ≤ 180µm
7.3 | Detector quantum efficiency of ≥ 70%
7.4 | Fully digital imaging chain & software
7.5 | Automatic gain & dose rate control
7.6 | Image rotation feature
7.7 | Integrated Dose rate Monitor with shortest warm-up time
7.8 | Matrix size of 2K x 2K pixel or better
7.9 | Multi Image review on dual 29” (or better) TFT flat panel monitor
7.1 | High image storage capacity
7.11 | Radiation warning light on display console
7.12 | Software such as DSA, Road mapping, edge enhancement, artifact reduction etc.
7.13 | Integrated DAP Monitor with shortest warm-up time
7.14 | Upgradeable to higher version
7.15 | PACS & DICOM interface for Image store, Query/Retrieve, Print, Work List Generation, MPPS, and archive
7.16 | Wireless DICOM compatibility
8 | Unit must be antimicrobial Protection to maintain high infection control standards
9 | Crystal-clear 2D image and high-resolution 2D reconstruction.Every tiny anatomical structures and implants become visible.
10 | Specify Weight and Dimensions of the unit
11 | Power consumption & Heat dissipation in BTU/h
12 | Quality control ,quality assurance and accaeptance test should be done by the company
13 | Safety assessment should be done by the company
14 | Software:
14.1 | Repeat /Reject Analysis software
14.2 | Statistical tool software
14.3 | Protocol analysis
15 | Accessories
15.1 | All necessary Accessories, consumables and disposables for starting-up.
15.2 | General Radiology QC tools for performance verification done by Enduser, if required.
15.3 | Lead aprons Quantity 5 ( double face 0.35 mm lead equivalen Light Weight Model
15.4 | Thyroid shield protection,Head Protection and Lead GogglesQuantity each 5 nos
15.5 | Sterile C-Arm Cover 50 Nos
16 | Trade-in (If relevant)
16.1 | Supplier should offer a Trade-in price for the following existing system:
16.2 | DESCRIPTION MANUFACTURER MODEL Sr.No YEAR OF PRUCHASE
16.3 | C-ARM X-RAY SIEMENS, GERMANY ARCADIS VARIAC 12781/647-90063888-XP9 21/07/2011
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
18 | Accessories
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
19 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at laest 3 years comprehensive from date of installation and equipment acceptance.

## Ultrasound
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| ULTRASOUND FOR RADIOLOGY | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Ultrasound machine and optional items.Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | Premium ultrasound machine capable of top-of-the-line and 4D imaging along with clear imaging for radiology
3 | Higher-end Ultrasound machine which includes all radiology application not limited to Abdominal, Vascular /Peripheral ,Obstetrical,Gynecological,Breast,SmallParts,Transcranial,Pediatric and Neonatal, Musculoskeletal,Urological,Cardiac etc.
4 | Should have four Active Probe slots.
5 | Transducer Types: • Convex Array Min 1-6 Mhz • Linear Array Min 3-10 Mhz • Matrix Array Min 6-15 Mhz • Small Linear probe Min L8-18 Mhz • Tranvaginal Prove (TVS) Min5-8 Mhz
6 | Contrast-Enhanced Ultrasound with TIC quantification
7 | Volume Navigation
8 | Must include Operating Modes • B-Mode • M-Mode • Color Flow Mode (CFM) • Extended Field of View • Power Doppler Imaging (PDI) • PW Doppler • CW Doppler • Volume Modes (3D/4D): • 3D Static • 4D Real Time • Anatomical M-Mode • Coded Contrast Imaging • Shear & strain Elastography •Onboard reporting •DVR •Scan Assist •Bsteer •AutoIMP •Thyroid & breast Productivity •Micro Bflow,Mesaure Assit Breast &OB,Compare Assit and flow QA.
9 | Unit should be fully supported to PACS system
10 | •Monitor Screen minimum 19-21′′ High-Resolution LCD •Min 10′′High Resolution Touch LCD screen for software menu
11 | specity Min/Max hight adjustment
12 | Horizontal/Vertical viewing angle of monitor
13 | Image Storage: •Storage Formats:DICOM – compressed/uncompressed, single/multiframe with/ without RawData •Export JPEG, WMV (MPEG 4), and AVI formats • support Storage Devices:USB Memory Stick,CD,DVD, Hard Drive stoarge
14 | Peripheral Options • User Interface Operator Keyboard • Digital B&W thermal printer • DVD video recorder • Digital color thermal printer • Foot Switch with programmable functionality •Gel warmer
15 | Connectivity & DICOM • Ethernet network connection• Verify• Print• Store• Modality Worklist• Storage Commitment• Query / Retrieve
16 | Having B/M color and PW/CW
17 | Unit Must haveHDMI /DVI/VGA ports
18 | Frame rate/Sampling rate must not be less then 150 frames/sec
19 | Must be supplied with 1 TB HDD or more
20 | Full DICOM and HL7 compatability
21 | Specify Dimensions and weight of the uint
22 | Unit should be provided with all necessary accessories and consumables.
23 | Must have universal power adapter 220-240v ,50/60HZ
Quality and Patient safety
22 | Unit must be FDA / CE marked .
23 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
24 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .