## Grossing station
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| GROSSING STATION | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Unit must be FDA approved/CE marked.
1 | Ergonomically designed to fit the working person’s needs (height, sitting or standing position) , Preferably Voice activated sensor which adjust to the Pre-defined end user .
2 | Must be Built with good quality standard Stainless steel.
3 | High Safety is Priority (Exhaust system, back draft/down draft/combination)
4 | Waste basket secured (trapdoor function and should connect to exhaust/filter)
5 | Formalin dispenser and formalin waste collection a must
6 | Convenience in using (work area, for keeping tools and/or samples)
7 | Easy operation (possible push buttons/touch control for easy dispensing/cleaning etc.)
8 | Grossing station should have “provision” for digital imaging (cameras/scanner) and dictation system both of which could be integrated and compatible with LIS.
9 | System must be having Electrical safety test IEC 6060-1
10 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
11 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
12 | Offered unit -History of incident,recalls and safety alerts .
13 | Can be integrated with other machines to provide specimen ID and tracking .
14 | Can also be use as biological safety cabinet .
15 | Offerd System must be capable to do the full documentation of the sample using the equipped multimedia devices .
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and histo lab technologist and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive 3 Years warrenty .

## Slide Stainer
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| AUTOMATED SLIDE STAINER | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Unit must be FDA approved/CE marked.
1 | Ergonomically designed (linear robotic type of stainer over centrifugal type)
2 | Built to last (if worst… can be easily converted to manual mode in cases of failure)
3 | Safety (with cover to act as fume hood, vapour aspiration and filters)
4 | Can load slides continuously and simultaneously.
5 | Easy operation (push buttons/touch control)
6 | Multiple/versatile programming of protocols.
7 | System must be having Electrical safety test IEC 6060-1
8 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
9 | Offered unit -History of incident,recalls and safety alerts .
10 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
11 | Offered unit -History of incident,recalls and safety alerts .
12 | Can be integrated with other machines- Preferably
13 | Load capacity should be more to accommodate more slides .
14 | Cover slipper must be included
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and histo lab technologist and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive 3 Years warrenty .

## IMH analyzer
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| FULLY AUTOMATED IMMUNO HISTO CHEMSTRY ANALYZER | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Unit must be FDA approved/CE marked.
1 | Ergonomically designed
2 | Safety (exhaust system and waste collection)
3 | Possible functions include de-waxing, antigen retrieval and staining
4 | Programs can be interrupted to allow addition of antibodies and/or slides even, if possible
5 | Must be an Open system
6 | Assistance in validating the stains (control blocks provided)
7 | System must be having Electrical safety test IEC 6060-1
8 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
9 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
10 | Offered unit -History of incident,recalls and safety alerts .
11 | Needs to specify the load capacity ..
12 | Software must be easy to use and programming must be user friendly
13 | Individual Temperature Controlled Slides.
14 | Regaent Preparation - Should be fully automated
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and histo lab technologist and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive 3 Years warrenty .