## OT Table
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To Be Filled
| Digital Mobile Advanced Protology Diagnostic System | 1 | Comply | Supporting Documents Reference | By Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | General Requirements:
3 | The system must be fully digital and mobile, designed for use in clinics.
4 | The system must be equipped with advanced imaging and diagnostic tools for proctology procedures such as a 1. Endoanal Manometry 2. Rectal Sensation Testing 3. Endoanal Ultrasound 4. Video Anoscopy
5 | The system should provide real-time imaging with clear, high-resolution outputs.
6 | The device must be easy to transport, with a mobile cart or case for easy relocation.
7 | The system must be compatible with various digital formats for image storage and transfer.
8 | Integrated laptop with minicamera for objective visualization during digital video anoscopy procedures.
9 | Self-illuminating integrated color LED light.
10 | Button-operable ultrasound probe for pelvic floor and sphincter muscle visualization.
11 | Probe frequency: Up to 14 MHz
12 | Capability for linear, angle, and volume measurements.
13 | Dedicated software for patient studies, including data charts and exam comparisons.
14 | Battery capacity: minimum 4 hours, with Bluetooth connectivity.
15 | Compact, lightweight, and easily transportable design.
16 | Short examination duration (approx. 10 minutes).
17 | Immediate availability of average pressure values during the exam.
18 | Automatic system activation without initial calibration
19 | Self-calibration for precise and continuous pressure measurements along the anal canal.
20 | Imaging and Visualization
20.1 | Display Type: High-definition, LED-backlit touch screen (minimum 21 inches).
20.2 | Resolution: At least 1920 x 1080 pixels for clear, high-quality images.
20.3 | Imaging Modes: Must support multiple imaging modes,
20.4 | Magnification: Digital zoom function for detailed imaging.
20.5 | Color Depth: Minimum of 24-bit color for accurate tissue differentiation.
22 | Endoscopic Functionality
22.1 | Diameter of Endoscope: Maximum diameter of 10mm, flexible design for patient comfort.
22.2 | Length: Minimum of 1 meter for ease of reach and manoeuvrability.
22.3 | Lighting: Built-in high-intensity LED light for clear visibility.
22.4 | Camera: Integrated high-resolution camera with adjustable focus for detailed visual examination.
23 | Unit must be antimicrobial Protection
24 | Specify Dimensions and weight of the units
25 | Unit should be provided with all necessary accessories and consumables.
26 | Must have universal power adapter 220-240v ,50/60HZ
4 | Quality and Patient safety
4.1 | Unit must be FDA / CE certified
4.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
4.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
4.4 | Quality control ,Quality assurance and acceptance test should be done by the company, QC tools for performance verification done by End-user, if required.
4.5 | Safety assessment should be done by the company
5 | Consumables
5.1 | Vendor must include prices of all standard and optional consumables.
6 | Available Additional Options
No. | Options Description |  |  |  | REQUIRED OR NOT REQUIRED
7 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | IEC 60601-2-38 Medical Electrical Equipment's part 2-38 Particular requirements for safety of Electrically Operated Hospital Beds
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment's recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty : Comprehensive at least 3 Years

## Sheet1
OT Table (UHS/MED/EQP/TENDER/004/2024)
S.No: | Local Distributor | Manufacturer | Models | Country of origin
1 | AL hayat | Schmitz | Diamond 60 BLK | Germany
2 | MPC | Medifa | Medifa 7000 | Germany
3 | Alphamed | Stryker | Operon 850 | USA
4 | EHS | Mizuho | MOT-VS5000 series /option-2-600 seres | Japan
5 | Fourmed | Hill-Rom -Trumpf | PST-500 | Germany
6 | Guld Drug | Steris | CMAX 3 | France-USA
7 | Leader Medical | Famed | optima SU-05.12 | Poland
8 | MPC | Maquet | Meera | Germany
9 | One Health | Mindray | Hybase V8 | PRC
10 | Quadri pharmaceutical | Merivaara | Grand Promerix | Euroupe/Finland
11 | Ultravision medical | ALVO Serenada-B | Serenada-B | Poland
12 | Virtus Trading | OPT Surgical system | VITA | Italy
13 | Westfort Trading | MAK | M0213 | Indonesia