## Radiology Ultrasound
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| ULTRASOUND FOR RADIOLOGY | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Ultrasound machine and optional items. Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | Premium ultrasound machine capable of top-of-the-line and 4D imaging along with clear imaging for radiology department.
3 | Higher-end Ultrasound machine which includes all OBG application not limited to Abdominal, Vascular /Peripheral, Obstetrics, Gynecology,Pediatrics,FetalCardioObstetrical,Gynecological,Breast,Small Parts,Transcranial,Pediatric and Neonatal, Musculoskeletal,Urological,Cardiac etc.
4 | Should have four Active Probe slots.
5 | Transducer Types: • Convex probe • Transvaginal Probe (TVS) • High frequency Linear array probe • Liner Hockey stick probe
6 | Volume Navigation
7 | Must include Operating Modes • B-Mode • Anatomical & ColorM-Mode • Colour Flow Mode (CFM) • Extended Field of View • Power Doppler Imaging (PDI) • PW Doppler • CW Doppler • 3D/4D, 4D Virtual HD • Auto Optimization • Colour Doppler • Compound Imaging • Contrast • DICOM, •Elastography •HD-Flow •HPRF • M-Mode Flow • Needle Visualization -(as optional item) •Power Doppler •PW Doppler •SonoAVC Follicle •SonoNT • SonoVCAD, •Speckle Reduction •STIC •Tissue Doppler • Tissue Harmonics •Tomographic •VOCAL •Volume Speckle Reduction •Coded Excitation/Harmonics and Uterine Trace
8 | Unit should be fully supported to PACS system
9 | •Monitor Screen minimum 21-24′′ High-Resolution LCD •Min 10′′High Resolution Touch LCD screen for software menu
10 | specify Min/Max Hight adjustment
11 | Horizontal/Vertical viewing angle of monitor
12 | Image Storage: •Storage Formats:DICOM – compressed/uncompressed, single/multiframe with/ without Raw Data •Export JPEG, WMV (MPEG 4), and AVI formats • support Storage Devices:USB Memory Stick,CD,DVD, Hard Drive storage
13 | Peripheral Options • User Interface Operator Keyboard • Digital B&W thermal printer • DVD video recorder -(as optional item) • Digital color thermal printer -(as optional item) • Foot Switch with programmable functionality •Gel warmer
14 | Connectivity & DICOM : • Ethernet network connection• Verify• Print• Store• Modality Worklist• Storage Commitment• Query / Retrieve
15 | Having B/M color and PW/CW
16 | Unit Must have HDMI /DVI/VGA ports
17 | Frame rate/Sampling rate must not be less then 150-250 frames/sec
18 | Must be supplied with 2 TB HDD
19 | Full DICOM and HL7 compatibility
20 | Specify Dimensions and weight of the unit
21 | Unit should be provided with all necessary accessories and consumables.
22 | Must have universal power adapter 220-240v ,50/60HZ
23 | Quality and Patient safety
24 | Unit must be FDA / CE marked .
25 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
26 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
27 | Accessories
43" LED TV with HDMI connection. 2.Room Light
28 | Consumables
Supplied with all standard consumables.
29 | Available Additional Option
No. | OPTIONS DESCRIPTION
30 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .

## HandHeld Ultrasound Machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To Be Filled
| Handheld Ultrasound machine | 1 | Comply | Supporting Documents Reference | By Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Handheld Ultrasound machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Specify Origin, Make and Model of the unit
3 | Offered model should not have a recall, and in case the offered model is purchased and there is a recall during warranty then the supplier has to provide a replacement within 5 days with the same features of the recalled model.
4 | The offered models of each part of the complete system should be from same manufacturer brand and Should have an authorized distributor in UAE.
5 | Portable, handheld, wireless ultrasound device designed exclusively for musculoskeletal (MSK) imaging and ultrasound-guided procedures.
6 | High-frequency linear probe: minimum 7–15 MHz for MSK structures.
7 | Minimum 150-element array.
8 | Penetration depth: up to 8–12 cm.
9 | Lightweight: ≤ 400 g.
10 | Suitable for imaging: -Muscles -Tendons -Ligaments -Bursae -Small joints -Peripheral nerves -Soft tissue lesions
11 | Imaging Modes:-B-Mode (2D) -M-Mode -Colour Doppler -Power Doppler -Needle visualization/needle enhancement -Tissue Harmonic Imaging -Speckle Reduction Imaging
12 | Interventional Procedure Features:
12.1 | Real-time needle guidance for injections and aspirations.
12.2 | Adjustable needle visualization and angle settings.
12.3 | Procedure presets for:
12.4 | Joint injections
12.6 | Bursa injections
12.7 | Peripheral nerve blocks
12.8 | Soft tissue aspirations
13 | Wireless operation with tablets/smartphones (iOS and Android).
14 | DICOM 3.0 support and PACS compatibility.
15 | Cloud-based image storage and sharing.
16 | Onboard storage for ≥ 500 images.
17 | Cine-loop recording and image annotation tools.
18 | Real-time display on tablet or smartphone.
19 | Adjustable gain, depth, focus, and TGC.
20 | Measurement tools: distance, area, and angle.
21 | Body-marking and annotation capabilities.
22 | Rechargeable lithium-ion battery.
23 | Minimum 3–4 hours continuous operation.
24 | USB-C or charging dock support.
25 | Fast-charging preferred.
26 | Real-time needle guidance for injections and aspirations.
27 | Adjustable needle visualization and angle settings.
28 | Quality and Patient safety
28.1 | Unit must be FDA / CE certified
28.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
28.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
28.4 | Quality control ,Quality assurance and acceptance test should be done by the company, QC tools for performance verification done by End-user, if required.
29 | Consumables
29.11
30 | Available Additional Options
No. | Options Description
31 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | IEC 60601-2-38 Medical Electrical Equipment's part 2-38 Particular requirements for safety of Electrically Operated Hospital Beds
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 10 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment's recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty : Comprehensive at least 3 Years