## Urodynamic machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Cardio-vascular Injector system for IR(Interventional radiology) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Cardio-vascular Injector and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | The offered system shall be a state-of-the-art angiographic contrast delivery/injector system designed for coronary angiography, cardiac catheterization, ventriculography,Neuro and peripheral vascular procedures
3 | Please specify make, model and origin of the machine.
4 | Main injector console/control unit
5 | Injector head assembly
6 | Touchscreen monitor/control panel
7 | Hand controller (AngioTouch or equivalent)
8 | Mobile cart with lockable wheels and bed rail mounting system
9 | Pressure monitoring interface
10 | X-ray synchronization interface cables
11 | Automatic air column detection sensor
12 | Contrast source monitoring sensor
13 | Syringe valve sensor
14 | Isolation manifold safety mechanism
15 | Automatic high/low pressure switching
16 | Real-time pressure monitoring
17 | Automatic injection stop in case of air detection
18 | Audible and visual alarms & Emergency stop function
19 | Contrast media reduction technology
20 | Contrast flow rate-Variable and fixed from 0.1 to 40 ml/sec or better
21 | Flow rate increment-0.1 ml/sec or better
22 | Saline flow rate-Approx. 1.6 ml/sec or better
23 | Injection volume range-0.1 to 99.9 ml or better
24 | Volume increment 0.1 ml
25 | User adjustable Pressure limit range from 200 to 1200 psi
26 | Refill rate Automatic/manual refill approx. 3 ml/sec
27 | LCA, RCA, LV/Ao, pigtail, selective, microcatheter, user-defined Injection modes
28 | Synchronize with major angiography systems-Siemens -icono Pro unit
29 | Hemodynamic integration
30 | Variable real-time flow control
31 | Hand-operated proportional controller
32 | unit must be available for demonstration
33 | Lightweight and compact design suitable for clinical rooms.
34 | Easy to clean surfaces and disinfectable parts.
35 | Unit must be antimicrobial protection
36 | Power supply: 220–240V, 50/60 Hz
37 | Please specify the weight and unit must be provided with good quality castors .
38 | Quality and Patient safety
38.1 | Device should be CE and FDA approved
38.2 | Offered System must comply with Electrical safety IEC 6060-1
38.3 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
38.4 | Electrical leakage and overload protection
39 | Accessories
39.1 | All necessary accessories
40 | Consumables
40.1 | Supplied with all standard consumables-Disposable patient kits, High-pressure tubing, Syringe assemblies, Manifold assemblies,Multi-procedure syringe usage capability Sterile single-use disposables
41 | Available Additional Option
No. | OPTIONS DESCRIPTION
42 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for doctors and other associated staff and Service training for Biomedical Engineering Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .