## Urodynamic machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Stationary EMG/NCS Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Urodynamic machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | The unit must include an advanced active noise cancellation system to reduce external electrical noise and interference during EMG and NCS testing.
3 | Please specify make, model, and origin of the machine.
4 | EMG Features:
4.1 | Electrode Inputs: 4 channels for surface EMG.
4.2 | Needle EMG Compatibility: Compatible with both surface and needle electrodes.
4.3 | Frequency Response: 20 Hz to 10 kHz or better.
4.4 | Amplification: High input impedance (>10 MΩ) to minimize signal distortion.
4.5 | Recording Mode: Capable of monopolar, bipolar, and common average reference electrode configurations.
4.6 | Real-time waveforms with adjustable gain, time base, and filter settings.
4.7 | Display and analysis of motor unit potentials and recruitment patterns.
4.8 | EMG Signal Processing: Integrated software for motor unit potential analysis, signal averaging, and automatic report generation.
5 | NCS Features:
5.1 | Nerve Conduction Velocity Measurement: Accurately measures motor and sensory nerve conduction velocity.
5.2 | Stimulation Types: Includes constant current stimulation, square wave, and biphasic waveforms.
5.3 | Recording Electrodes: Compatible with surface electrodes for sensory and motor nerve conduction studies.
5.4 | Conduction Study Parameters: Ability to measure latencies, amplitudes, and conduction velocities for motor and sensory studies.
5.5 | Distance Measurement: Automatic distance calculation for nerve conduction study accuracy
6 | High-resolution colour touchscreen or monitor with a minimum size of 21 to24 inches.
8 | Intuitive, user-friendly interface with customizable settings.
9 | Data Visualization: Clear, high-quality graphical representations of EMG and NCS results.
10 | Data Exporting: Ability to export results in standard formats (e.g., PDF, CSV, DICOM).
11 | Diagnostic Software: Includes built-in diagnostic software for both EMG,NCS ,EP and autonomic. And can be connected to Doctor's PC to review the report.
12 | Sufficient internal storage 2 TB to hold patient data, test results, and historical records.
13 | Support for integration with hospital or clinic networks (via HL7, PACS).
14 | Multiple user profiles for different operators with password protection for patient data security.
15 | Automatic generation of standardized reports with customizable templates.
13 | Ergonomics & Portability
13.1 | Mobile cart with lockable wheels.
13.2 | Lightweight and compact design suitable for clinical rooms.
14 | Easy to clean surfaces and disinfectable parts.
15 | Unit must be antimicrobial protection
16 | Power supply: 220–240V, 50/60 Hz
17 | Quality and Patient safety
17.1 | Device should be CE and FDA approved
17.2 | Please specify the weight and unit must be provided with good quality castors .
17.3 | Offered System must comply with Electrical safety IEC 6060-1
17.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
18 | Accessories
18.1 | All necessary accessories, including Electrodes (surface and needle types).
19 | Consumables
19.1 | Supplied with all standard consumables
20 | Available Additional Option
No. | OPTIONS DESCRIPTION
21 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Neurology technicians and doctors and other associated staff and Service training for Biomedical Engineering Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .