## Urodynamic machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Urodynamic machine (Urology Dept) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Urodynamic machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Please specify make, model, and origin of the machine.
3 | System to be used in Urology clinic for non-invasive screening to measure the change in urine flow rate within a unit time during natural urination by means of mass transducer.
4 | Multi-channel urodynamic system for cystometry, uroflowmetry, pressure-flow studies, and EMG monitoring.
5 | Minimum 3 pressure channels (intravesical, abdominal, detrusor) and optional EMG channels.
6 | Uroflowmetry: Integrated uroflow measurement with real-time graphical display.
6.1 | lightweight and easy to clean commode chair
6.2 | High-resolution color touchscreen monitor.
6.3 | software: Advanced urodynamic analysis software with:
7 | Real-time graphing and pressure tracing.
8 | Automatic report generation.
9 | Data export in standard formats (PDF, DICOM, CSV).
10 | Connectivity: USB/wireless and network-enabled for HIS integration
11 | Patient Safety: Non-invasive and minimally invasive options; CE/FDA-approved consumables.
12 | Measurement Accuracy
12.1 | Pressure measurement accuracy: ±1 cmH2O.
12.2 | Flow rate accuracy: ±5% or better.
12.3 | EMG signal accuracy: High-fidelity, low-noise s
13 | Ergonomics & Portability
14 | Mobile cart with lockable wheels.
15 | Lightweight and compact design suitable for clinical rooms.
16 | Easy to clean surfaces and disinfectable parts.
17 | Unit must be antomicrobial protection
18 | Power supply: 220–240V, 50/60 Hz
19 | Quality and Patient safety
19.1 | Device should be CE and FDA approved
19.2 | Please specify the weight and unit must be provided with good quality castors .
19.3 | Offered System must comply with Electrical safety IEC 6060-1
19.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
20 | Accessories
20.1 | All necessary accessories, including catheters, pressure transducers, software, and patient interface devices.
21 | Consumables
21.1 | Supplied with all standard consumables-1)Single-use catheters of various sizes.2)Transducer kits and calibration solutions.3)EMG electrodes (disposable and reusable options).
22 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
23 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .