## Coronary Microvascular Dysfunct
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Coronary Microvascular Dysfunction (CMD) Diagnostic unit | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Please specify make, model, and origin of the machine.
3 | The proposed tender aims to procure a dedicated diagnostic unit capable of reliably performing IMR and CFR measurements in a catheterization laboratory setting.
4 | Pressure and temperature-sensing guidewires for invasive IMR measurement.
5 | Compatible sensors for CFR assessment (thermodilution or Doppler-based).
6 | Catheterization lab-compatible monitor with real-time display of coronary physiology parameters.
7 | Integration capability with existing hemodynamic monitoring systems and imaging consoles.
8 | Consumables (wires, catheters, calibration tools) for at least 100 procedures.
9 | Real-time calculation of IMR and CFR with visual display.
10 | Data logging and export to electronic medical record (EMR) systems.
11 | Analysis tools for patient follow-up and research reporting.
12 | Alarm/indicator system for measurement errors or poor signal quality.
13 | IMR and CFR measurement accuracy: ±10% or better.
14 | Measurement feasible in vessels ≥2 mm in diameter.
15 | Compatible with hyperemic agents (adenosine, regadenoson, or equivalent).
16 | Safe for use in adult patients undergoing coronary angiography.
17 | Please specify the weight and unit must be provided with good quality castors .
18 | Minimal additional procedure time (<15 minutes per measurement).
24 | The device must provide clinical-grade accuracy validated by peer-reviewed studies and real-world performance data
25 | Unit must be antomicrobial protection
26 | Power supply: 220–240V, 50/60 Hz
27 | Quality and Patient safety
27.1 | Device should be CE and FDA approved
27.3 | Offered System must comply with Electrical safety IEC 6060-1
27.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
28 | Accessories
Supplied with all standard accessories
29 | Consumables
Supplied with all standard consumables
30 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
31 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Optical Coherence Tomography (O
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Optical Coherence Tomography (OCT) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
0 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Please specify make, model, and origin of the machine.
2 | Provides actionable OCT and angiographic insights, enabling greater efficiency during PCI procedures and supporting improved clinical outcomes.
3 | Type: Frequency-domain / Fourier-domain OCT
4 | Resolution:Axial: 10–20 μm,Lateral: 20–40 μm
5 | Imaging Depth: Up to 10 mm in coronary vessels
6 | Frame Rate: 100–180 frames per second (real-time imaging)
7 | Catheter Compatibility: 5 Fr–6 Fr guide catheters
9 | Automated plaque characterization (fibrous, lipid-rich, calcified)
10 | Stent apposition and expansion assessment
11 | Vessel diameter and lumen area measurements
12 | Lesion length calculation and optimal stent sizing
13 | Automated detection of thrombus and dissections
14 | Angiographic Integration
14.1 | Real-time co-registration with angiography for precise lesion localization
14.2 | Fusion of OCT and angiographic imaging for enhanced procedural guidance
14.3 | Supports roadmap overlay for stent placement
14.4 | Minimal additional procedure time (<15 minutes per measurement).
15 | Procedural Efficiency Features:
15.1 | AI-powered automated measurements and workflow optimization
15.2 | Rapid pullback imaging (≤3 seconds per coronary segment)
15.3 | Reduced contrast and procedure time
15.4 | Data export for post-procedural analysis and reporting
16 | Software & AI Capabilities
16.1 | AI-assisted detection of high-risk plaque features
16.2 | Automated stent selection recommendations
16.3 | Interactive 3D vessel reconstruction
16.4 | Integration with hospital PACS and EMR system
17 | Unit must be antomicrobial protection
18 | Power supply: 220–240V, 50/60 Hz
19 | Quality and Patient safety
19.1 | Device should be CE and FDA approved
19.2 | Offered System must comply with Electrical safety IEC 6060-1
19.3 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
20 | Accessories
Supplied with all standard accessories
21 | Consumables
Supplied with all standard consumables
22 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
33 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .