## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Delivery Bed | 6 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
The technical offer must clearly state the standards and specifications of the proposed Labor and Delivery (L&D) Bed designed for obstetric use including labor, delivery, and recovery. All specifications shall be detailed clearly and supported by relevant manufacturer reference documents and brochures.
1 | Bed Structure & Design
1 | The bed shall be ergonomic and designed for obstetric procedures.
2 | The bed shall be electrically operated for all primary functions including height, backrest, and positioning. The leg/foot section shall be of stow-and-go design, allowing it to be stored within the bed frame and deployed quickly for delivery without requiring complete removal.
3 | The frame shall be made of high-quality, corrosion-resistant material.
4 | The mattress platform shall be multi-section (minimum 3–4 sections).
5 | Mattress shall be seamless or covered to prevent fluid penetration, waterproof, antibacterial, and easy to clean.
6 | Include ergonomic hand grips for the mother during labor.
7 | The bed shall allow easy and quick conversion from labor position to delivery position without the use of tools.
8 | Include removable stainless steel delivery collection tray or any drainage collection system.
9 | The bed shall have electric adjustment for: • Height adjustment: minimum range of approximately 55 - 95 cm (±5 cm) • Backrest adjustment: minimum 0° to 65° (±5°) • Trendelenburg / Reverse Trendelenburg positions: minimum 12° (±2°)
10 | The bed shall include hand control for patient and caregiver.
11 | The system shall include emergency CPR quick release function.
12 | The bed shall include adjustable lithotomy leg supports with padding.
13 | The bed shall provide unobstructed and ergonomic access to the perineal area during delivery.
14 | The bed shall allow chair and upright labor positions.
2 | Mobility and Mechanical Design
1 | The bed shall be mounted on central braking system with lockable castors.
2 | The bed shall include steering mechanism for easy transport.
3 | The bed shall support minimum patient weight capacity of 200–250 kg.
3 | Accessories and Safety Features
1 | Specify Dimensions and weight of the unit.
2 | Specify Min/Max Hight adjustment.
3 | The bed shall include side safety rails with easy locking mechanism.
4 | The bed shall allow rapid repositioning during emergencies.
5 | The bed shall include an adjustable IV pole with multiple hooks.
6 | Unit should be provided with all necessary accessories and consumables.
7 | The system shall include electrical overload protection to prevent damage to motors and ensure safe operation under excessive load conditions.
8 | The bed shall include a rechargeable battery backup system capable of supporting a minimum of 20–30 full adjustment cycles during power failure.
9 | Must have universal power adapter 220-240v ,50/60HZ
4 | Quality and Patient safety
1 | Unit must be FDA / CE marked .
2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
3 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
5 | Accessories
Should have All standard Accessories must be included as per manufacturer .
6 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
2
7 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training ifor operator and Physians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation