## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Dermatology Laser Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must clearly specify the standards and detailed specifications of the proposed Laser System intended for use in Dermatology and Plastic Surgery for multi-application treatments. All technical details, including laser type, wavelength(s), energy output, cooling system, safety features, and clinical applications, must be clearly stated. All requested specifications must be supported with official reference documents
1 | Laser Technology
Specify: -Manufacturer / make -Model -Country of origin -Laser type -Laser class -Maximum power output in watts -Power requirement and heat dissipation
The system shall support treatments including: -Hair removal -Vascular lesions -Pigmented lesions -Skin rejuvenation
The system shall be dual-wavelength laser platform : Alexandrite (~755 nm) AND Nd:YAG (~1064 nm) OR equivalent clinically proven wavelengths
The system shall allow selection between wavelengths without changing device.
2 | Energy & Performance
Wide adjustable fluence range suitable for all skin types (I–VI).
Adjustable pulse duration (short + long pulse capability).
Adjustable repetition rate (Minimum 10 Hz or higher preferred).
High peak power output for effective treatment.
Uniform energy distribution across spot.
3 | Spot Size & Applicators
Multiple spot/window sizes shall be available ranging from 2 mm to ≥25 mm
Automatic spot size recognition preferred.
Handpieces shall be lightweight and ergonomic.
Fast switching between spot sizes preferred.
4 | Cooling System
Integrated skin cooling system Acceptable types: -Cryogen spray cooling, specify type of cryogas used and average consumption per treatment/session -Contact cooling -Air cooling (if clinically effective)
Cooling shall be adjustable and synchronized with laser pulse.
Supplier shall clearly specify all operational consumables and their expected replacement frequency and running cost
5 | User Interface
Touchscreen display (minimum ~8–10 inch preferred).
Pre-set treatment protocols available.
Display shall show: -Fluence -Pulse duration -Frequency -Spot size
System shall allow manual parameter adjustment.
6 | Design, Accessories and Safety Features
Emergency stop button (E-stop).
Key switch access control.
Footswitch activation.
The laser handpiece shall include a detachable spacer to ensure accurate treatment distance and patient safety.
Spacer attachments of different sizes shall be supplied for various applications
Visual and audible emission warning.
System fault detection with auto shutdown.
User access levels / password protection.
Mobile unit with lockable wheels.
The system shall include: -All required handpieces -Footswitch -Protective goggles (operator + patient, wavelength-specific)
Specify Dimensions and weight of the unit
Specify the expected flash lamp life in number of shots/pulses, recommended replacement interval and cost of the flash lamp
Proper hazard labeling as per ANSI
System shall include internal protection from power fluctuations.
Unit should be provided with all necessary accessories and consumables.
7 | Quality and Patient safety
Unit must be FDA / CE marked .
Unit must comply ISO 13485 Quality management system for medical device or equivalent .
The offered laser system shall comply with IEC 60825-1 laser safety standards & ANSI Z136
System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
8 | Accessories
Should have All standard Accessories must be included as per manufacturer .
9 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
2
10 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training ifor operator and Physians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation