## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Neonatal Resuscitaire for L&D | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must clearly specify the standards and specifications of the proposed Neonatal Resuscitaire intended for use in L&D for immediate newborn stabilization and resuscitation .Please specify the details clearly and all the requested specs info must be supported with reference documents.
1 | Radiant Heating System
Radiant heater with focused heating technology for uniform heat distribution.
The system shall include servo-controlled and manual temperature control modes.
Temperature setting range approximately 34.5°C – 37.5°C.
The unit shall include pre-warm mode to prepare the warmer before delivery.
The unit shall include a digital display for temperature and system parameters.
The heater head shall rotate or tilt away to allow full access to the baby during procedures.
The heater shall include an automatic safety mechanism that reduces or stops heater output when the heater is not positioned correctly above the infant.
Skin temperature monitoring using neonatal skin probe.
2 | Neonatal Resuscitation System
The unit shall include an integrated neonatal resuscitation module.
The system shall include an integrated Neopuff T-Piece Resuscitator module or equivalent.
The resuscitation system shall support Positive Pressure Ventilation (PPV).
The system shall provide adjustable Peak Inspiratory Pressure (PIP) and Positive End Expiratory Pressure (PEEP).
The resuscitation system shall include safety pressure relief valve.
The system shall include pressure manometer display for airway pressure monitoring.
3 | Gas Supply and Flow Control
The system shall include oxygen / air blender with adjustable FiO₂ range from 21% to 100%.
The system shall include flow meter suitable for neonatal resuscitation.
The system shall support connection to hospital medical gas supply (Oxygen and Medical Air).
The unit shall also allow operation with gas cylinders when pipeline supply is unavailable.
The unit shall include gas cylinder holder
4 | Patient Platform and Bassinet
The patient platform shall include tilting mattress platform.
Mattress shall be radiolucent and easy to clean.
The bassinet shall include transparent protective side panels for neonatal safety.
The system shall include integrated X-ray cassette tray below the mattress, allowing imaging without moving the newborn.
5 | Weighing System
The unit shall include an integrated neonatal weighing scale built into the patient platform or mattress tray.
The scale shall include digital weight display visible to the operator.
The scale shall include tare / zero function to allow weighing with blankets or accessories.
The weighing scale shall have minimum capacity of 10 kg.
The scale accuracy shall be ±10 g or better.
6 | Design, Accessories and Safety Features
The unit shall include LED examination light for procedures.
The unit shall include APGAR timer.
The system shall include electric height adjustment (elevator base) for ergonomic use.
The system shall be mounted on mobile base with lockable castor wheels.
The unit shall include storage drawers or shelves for accessories.
The unit shall include neonatal suction system or provision for connection to external suction.
The system shall provide audible and visual alarms, including: • High / low temperature • Skin probe failure • Heater malfunction • Power failure
The unit shall be suitable for term and preterm neonates immediately after birth.
Specify Dimensions and weight of the unit
Specify Min/Max Hight adjustment
Unit should be provided with all necessary accessories and consumables.
Must have universal power adapter 220-240v ,50/60HZ
7 | Quality and Patient safety
Unit must be FDA / CE marked .
Unit must comply ISO 13485 Quality management system for medical device or equivalent .
System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
8 | Accessories
Should have All standard Accessories must be included as per manufacturer .
9 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
2
10 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training ifor operator and Physians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation