## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| IVF OT CEILING LIGHT | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
The technical offer must clearly state the standards and specifications of the proposed ceiling-mounted LED operating light designed for use in an In Vitro Fertilization (IVF) procedure room. The system shall provide high-quality illumination for oocyte retrieval, embryo transfer, and other gynecological procedures. All specifications shall be clearly detailed and supported by relevant manufacturer reference documents and brochures
1 | General Requirements
The light shall provide bright, homogeneous, and shadow-free illumination for delicate gynecological procedures
The system shall utilize long-life LED technology with minimal maintenance requirements
The vendor shall be responsible for conducting the pre-installation site assessment and confirming site readiness prior to installation
The vendor shall be responsible for dismantling and removing the existing operating light
The vendor shall restore the installation site to its original condition after completion of installation
The vendor shall provide all necessary materials, accessories, and labor required for complete installation and commissioning of the new operating light.
Trade-in of the existing operating light is preferred, where applicable. Existing Operating Light: Model: AMICO MIRA 65 Configuration: Ceiling Mounted – Single Head Asset No: MEQ4516
2 | Design & Construction
The operating light shall be ceiling-mounted LED type suitable for IVF procedure rooms
Single-arm suspension system
Balanced suspension arm allowing effortless positioning while maintaining the selected position without drift
The operating light shall be manufactured from high-quality, durable materials resistant to impact, corrosion, and routine cleaning/disinfection
Sterilizable removable handle for repositioning while maintaining the sterile field
3 | Illumination Performance
Adjustable light intensity, light field diameter (focus size), and color temperature
Illumination intensity shall be adjustable within a range of 65,000 to 75,000 lux
The vendor shall specify the working distance at which the stated illumination (65,000–75,000 lux) is achieved
High Color Rendering Index (CRI) to ensure accurate tissue color recognition
Low heat emission to maintain a comfortable procedure environment
Minimal glare to reduce operator eye fatigue
Uniform light field with excellent depth of illumination
4 | Clinical Features
Suitable for oocyte retrieval, embryo transfer, and minor gynecological procedures
Suitable for prolonged procedures without significant heat generation
Excellent shadow management to maintain illumination when clinicians are working around the surgical field
5 | Accessories and Safety Features
Specify overall dimensions and weight of the unit
Specify the minimum and maximum height adjustment range
Easy-to-clean sealed surface suitable for infection control
Low power consumption with energy-efficient LED technology
Compatible with 220–240 V, 50/60 Hz electrical supply
6 | Quality and Patient safety
Unit must be FDA / CE marked .
Unit must comply ISO 13485 Quality management system for medical device or equivalent .
System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
7 | Accessories
All standard accessories supplied by the manufacturer shall be included
8 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
2
9 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End User Training for operators and physicians is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be minimum 3 years comprehensive from date of installation