## Integrated arthroscopy pump
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Integrated arthroscopy pump | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Integrated arthroscopy pump and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Please specify make, model and origin of the machine.
3 | The offered system shall be a microprocessor-controlled integrated arthroscopy pump for fluid distension, irrigation, and aspiration during arthroscopic procedures.
4 | The system shall be suitable for knee, shoulder, hip, elbow, ankle, wrist, and small joint arthroscopy.
5 | The system shall provide integrated inflow and outflow fluid management.
6 | The system shall be capable of interfacing with an arthroscopic shaver/resection console for synchronized operation.
7 | Pump Performance
7.1 | Dual-channel peristaltic pump system (inflow and outflow).
7.2 | Microprocessor-controlled pressure regulation.
7.3 | Continuous real-time intra-articular pressure monitoring.
7.4 | Automatic pressure compensation during instrument activation.
7.5 | Adjustable irrigation pressure.
7.6 | Adjustable flow rate.
7.7 | Adjustable suction level.
7.8 | Stable joint distension with rapid response to pressure fluctuations.
7.9 | Automatic priming function, Automatic leak detection, Automatic air removal function, Automatic self-test during startup.
8 | User Interface(Display )
8.1 | Large colour touchscreen display.
8.2 | Intuitive graphical user interface.
8.3 | Real-time display of: Pressure rate ,Flow rate,Suction,System status, Alarm messages
9 | Fluid Management
9.1 | Precise pressure regulation, Adjustable irrigation flow, Adjustable aspiration flow, Inflow-only mode, Combined inflow/outflow mode, Combined inflow/outflow mode, Automatic pressure maintenance, Cassette-based tubing system for quick setup, Disposable sterile tubing, Gravity irrigation option. sets.
10 | Safety Features
10.1 | Over-pressure protection.
10.2 | Low-pressure alarm.
10.3 | Occlusion detection.
10.4 | Air-in-line detection.
10.5 | Empty fluid bag alarm.
10.6 | Waste container full alarm.
10.7 | Audible and visual alarm indicators.
10.8 | Emergency stop function.
10.9 | Automatic shutdown under fault conditions.
11 | unit must be available for demonstration
12 | Easy to clean surfaces and disinfect able parts.
13 | Unit must be antimicrobial protection
14 | Power supply: 220–240V, 50/60 Hz
15 | Please specify the weight and unit Heat dissipation
16 | Quality and Patient safety
16.1 | Device should be CE and FDA approved
16.2 | Offered System must comply with Electrical safety IEC 6060-1
16.3 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
16.4 | Electrical leakage and overload protection
17 | Accessories
17.1 | All necessary accessories
18 | Consumables
18.1 | Supplied with all standard consumables-Disposables
19 | Available Additional Option
No. | OPTIONS DESCRIPTION
20 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Ortho doctors and other associated staff and Service training for Biomedical Engineering Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .

## Tenex Ultrasound Unit(Ultrasoni
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Tenex Ultrasound Unit(Ultrasonic tenotomy ) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Tenex ultrasound machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Please specify make, model and origin of the machine.
3 | The offered system shall be a state-of-the-art percutaneous ultrasonic tenotomy system intended for minimally invasive treatment of chronic tendinopathy, calcific tendinitis, plantar fasciitis, scar tissue, and selected bone lesions under ultrasound guidance.
4 | The system shall be suitable for treatment of:Tennis elbow (Lateral Epicondylitis) Golfer's elbow Achilles tendinopathy Patellar tendinopathy Plantar fasciitis Rotator cuff tendinopathy Gluteal tendinopathy Hamstring tendinopathy Calcific tendinitis Bone spur/Haglund deformity (where applicable) Other chronic degenerative tendon disorders under ultrasound guidance.
5 | The system shall utilize ultrasonic energy to fragment, emulsify, irrigate, and aspirate pathological soft tissue and calcific deposits through a single-use sterile handpiece.
6 | Compact mobile console with integrated touchscreen user interface.
7 | Colour touchscreen display for system operation and procedure settings.
8 | Adjustable ultrasonic power settings.
9 | Integrated irrigation and aspiration control.
10 | Integrated peristaltic aspiration pump.
11 | Automatic priming and system self-test.
12 | Audible and visual alarms.
13 | Footswitch operated.
14 | Ultrasonic Handpiece(Probe)
14.1 | Sterile, single-use disposable ultrasonic Microtip/handpiece.
14.2 | Integrated irrigation and aspiration tubing.
14.3 | Ergonomic lightweight design.
14.4 | Multiple tip configurations suitable for different anatomical locations.
14.5 | Suitable for both soft tissue and calcific lesion removal.
14.6 | Optional bone-specific Microtip shall be available.
15 | Irrigation and Aspiration System
15.1 | Pressure-controlled irrigation system.
15.2 | Adjustable aspiration flow.
15.3 | Integrated waste collection system.
15.4 | Automatic fluid management.
16 | unit must be available for demonstration
17 | Easy to clean surfaces and disinfect able parts.
18 | Unit must be antimicrobial protection
19 | Power supply: 220–240V, 50/60 Hz
20 | Please specify the weight and unit must be provided with good quality castors .
21 | Quality and Patient safety
21.1 | Device should be CE and FDA approved
21.2 | Offered System must comply with Electrical safety IEC 6060-1
21.3 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
21.4 | Electrical leakage and overload protection
22 | Accessories
22.1 | All necessary accessories
23 | Consumables
23.1 | Supplied with all standard consumables-Disposable sterile Microtips, Irrigation tubing sets, Aspiration tubing, Procedure packs, Waste collection accessories
24 | Available Additional Option
No. | OPTIONS DESCRIPTION
25 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Ortho doctors and other associated staff and Service training for Biomedical Engineering Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .