## Hydrotherapy Pool
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Hydrotherapy Pool with Treadmill, Cycle, and Hoist | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Please specify make, model, and origin of the machine.
1.1 | Please specify unit dimensions.
1.2 | Pool should be static depth or dynamic depth (preferred dynamic depth. Needs site inspection to determine possibility.)
1.3 | Depth range should be up to 1.5m.
1.4 | Dynamic depth should be remote controlled.
1.5 | Unit should have chair for seated shoulder mobility exercises.
1.6 | Pool should be fully wheelchair accessible.
1.7 | Pool should have stainless steel handlebars for patient safety.
1.8 | Water temperature should be adjustable and accommodate temperature range of at least 20 - 40 degrees C.
1.9 | Unit should have temperature control alarm to prevent overheating the water during malfunction.
1.1 | Water currents preffered to be paddle wheel or propeller.
1.11 | Water currents should be adjustable with controls.
1.12 | Unit should come with various pool accessories such as noodles and floats.
1.13 | Unit should come with storage rack or container for all pool accessories.
1.14 | Unit should come with emergency evacuation equipment and life saving equipment.
1.15 | Site limitation: Clear depth from finish floor until footing pcc top: 1.5m
2 | Underwater Treadmill | 1
21 | Please specify make, model, and origin of the machine.
2.2 | Please specify unit dimensions.
2.3 | Unit should be motorized with controllable speed.
2.4 | Unit should have emergency stop button.
2.5 | FDA approved / CE Marked.
3 | Underwater Cycle | 1
3.1 | Please specify make, model, and origin of the machine.
3.2 | Please specify unit dimensions.
3.3 | Unit should be motorized with controllable speed.
3.4 | Unit should have handle bars for stability.
3.5 | FDA approved / CE Marked.
4 | Patient Hoist | 1
4.1 | Please specify make, model, and origin of the machine.
4.2 | Please specify unit dimensions.
4.3 | Hoist should be able to be controlled remotely.
4.4 | Patient hoist should have max patient weight of at least 200kg.
4.5 | FDA approved / CE Marked.
4.6 | Offered System must comply with Electrical safety IEC 6060-1.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .