## Other Machines
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Unweighting Gait-Training System with Parallel Bars | 2 | Comply | Supporting Documents Ref. | by Hospital
2 | Paraffin Wax Bath Tub | 2
3 | Magnetotherapy Machine | 1
4 | Red Cord Suspension System | 1
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Unweighting Gait-Training System with Parallel Bars | 2
Please specify make, model, and origin of the machine.
Please specify any site requirements if needed.
1.1 | Unit should be ceiling-mounted.
1.2 | Unit should feature clinical partial body-weight-support/unweighting gait-training system for overground gait rehabilitation
1.3 | Adjustable partial unloading of the patient's body weight during gait training
1.4 | Support mechanism: Motorized or mechanically controlled body-weight support system providing stable and controlled vertical support
1.5 | Height adjustment: Powered/electric height adjustment preferred; sufficient vertical range to accommodate pediatric/adult patients as specified
1.6 | Adjustable pelvic/thoracic gait-training harness suitable for prolonged therapy sessions
1.7 | Unit includes minimum 2–3 harness sizes, or equivalent adjustment range covering various patient sizes.
1.8 | Pelvic movement: Suspension should permit appropriate anterior/posterior and lateral pelvic movement required for natural gait training, while maintaining patient safety
1.9 | Integrated fall-arrest/safety mechanism independent of normal weight-support operation
1.1 | Emergency stop accessible to therapist and/or patient where appropriate
1.11 | Integrated bilateral parallel bars suitable for gait training and balance exercises
1.12 | Height adjustable parallel bar height to accommodate different patient heights and clinical requirements
1.13 | Load monitoring: Real-time measurement/display of supported load and/or percentage body-weight unloading
1.14 | Controls shall allow therapist to increase/decrease body-weight support during treatment
1.15 | FDA approved / CE Marked .
1.16 | Offered System must comply with Electrical safety IEC 6060-1.
1.17 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Paraffin Wax Bath Tub | 1
Please specify make, model, and origin of the machine.
2.1 | Operating Temperature: 52°C – 57°C.
2.2 | Small sized wax tub.
2.3 | FDA approved / CE Marked .
3 | Magnetotherapy Machine | 1
Please specify make, model, and origin of the machine.
3.1 | Unit must include 2 applicators + Chair for pelvic and lower back stimulation.
3.2 | Output of up to 3 tesla.
3.3 | Frequency 1-160 Hz.
3.4 | Magnetic output can support up to 6 Different waveforms.
3.5 | Unit must come with large LCD color touchscreen.
3.6 | Variety of preset programmes & possibility to customize or create own programmes.
3.7 | FDA approved / CE Marked.
3.8 | Offered System must comply with Electrical safety IEC 6060-1.
3.9 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
4 | Red Cord Suspension System | 1
Please specify make, model, and origin of the machine.
Please specify any site requirements if needed.
4.1 | Professional clinical suspension therapy system suitable for physiotherapy, rehabilitation and neuromuscular training.
4.2 | Multiple independently adjustable suspension points allowing simultaneous suspension of different body segments.
4.3 | Height adjustment of individual suspension ropes/straps to accommodate different patient positions and exercise configurations
4.4 | Sliding transverse/adjustable suspension system allowing suspension points to be repositioned along the frame
4.5 | Ropes should be high-strength, durable suspension ropes suitable for repeated clinical use
4.6 | Secure locking mechanism preventing unintended movement when the system is loaded
4.7 | Unit must come with multiple interchangeable sling configurations for cervical, upper-limb, pelvic, trunk and lower-limb support
4.8 | Maximum working load of at least 150 kg.
4.9 | Complete accessory set required for normal clinical operation should be included.
4.10 | FDA approved / CE Marked.
4.11 | Offered System must comply with Electrical safety IEC 6060-1.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .

## Paediatric Gym Equipment
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Paediatric Upper and Lower Limb Exerciser | 1 | Comply | Supporting Documents Ref. | by Hospital
2 | Paediatric Treadmill | 1
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Paediatric Upper and Lower Limb Exerciser
Please specify make, model, and origin of the machine.
1.1 | Unit should have arms and legs training simultaneously.
1.2 | Unit should have active, passive,and resistance treatment modes.
1.3 | Unit should come with a color touch screen to adjust treatment parameters.
1.4 | Unit should have an adjustable height.
1.5 | Unit upper and lower trainer radius should be adjustable.
1.6 | Unit should be easy to be relocated (wheels).
1.7 | Unit should be silent during operation.
1.8 | FDA approved / CE Marked .
1.9 | Offered System must comply with Electrical safety IEC 6060-1.
1.1 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Paediatric Treadmill
Please specify make, model, and origin of the machine.
2.1 | Unit should have adjustable speed.
2.2 | Unit should have adjustable safety handlebars.
2.3 | Unit should have adjustable incline angle.
2.4 | Unit should have an emergency stop button.
2.5 | FDA approved / CE Marked.
2.6 | Offered System must comply with Electrical safety IEC 6060-1.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .

## Small Items
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Exercise Mat | 4 | Comply | Supporting Documents Ref. | by Hospital
2 | Resistive Exercise Band and Ball set | 2
3 | Wristlet Weight set | 2
4 | Ankle Weight set | 1
5 | Shoulder Wheel | 1
6 | Weight Rack Tree | 2
7 | Gym Organiser | 2
8 | Balance Training Items | 2
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Exercise Mat | 4
1.1 | Unit should be at least 96cm in width and 180cm in length.
1.2 | Unit should be made of durable material.
2 | Resistive Exercise Band and Ball set | 2
2.1 | One set should come with bands and balls in a range of resistances.
2.2 | Units should be made of durable material.
3 | Wristlet Weight set | 2
3.1 | One set should come with at least 0.5kg, 1kg, and 2kg.
3.2 | Each unit should be one-size-fits-all to accommodate for all patient types.
4 | Ankle Weight set | 1
4.1 | One set should come with at least 0.5kg, 1kg, and 2kg.
4.2 | Each unit should be one-size-fits-all to accommodate for all patient types.
5 | Shoulder Wheel | 1
Please specify make, model, and origin of the machine.
5.1 | Unit should be wall-mounted.
5.2 | Unit must allow for bi-directional movement.
5.3 | Unit should be height adjustable.
5.4 | Unit should have adjustable resistance for progressive exercises.
5.5 | Unit should be able to adjust arm length.
5.6 | FDA approved / CE Marked.
6 | Weight Rack Tree | 2
6.1 | Unit should include a range of free weights from 0.5 kg and above.
6.2 | Unit should be of rubber material.
7 | Gym Organiser rack | 2
7.1 | Unit should include adequate shelves and space to accommodate exercise mats, exercise bads, ankle + wrist weight sets, etc.
8 | Balance Training Items | 2
8.1 | Proposed unit should include wobble exercise board, bosu ball balance trainer, stability discs.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .

## Sensory Room
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Paediatric Sensory Motor Integration Room + Occupational Therapy Items | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Please specify make, model, and origin of the machine.
1 | The proposed system shall be suitable for use within a hospital rehabilitation/physiotherapy and occupational therapy environment and shall support therapeutic activities involving: • Vestibular processing • Proprioceptive input • Tactile processing • Motor planning and coordination • Balance and postural control • Gross motor development • Fine motor activities • Visual-perceptual activities • Bilateral coordination • Functional activities of daily living • Regulation/calming activities
2 | The Contractor shall provide a complete and operational installation, including all necessary brackets, fixings, ropes, carabiners/connectors, adjustment mechanisms, protective padding and other accessories, whether or not each accessory is individually listed.
3 | Proposed system should be suitable for continuous use in a healthcare/rehabilitation environment.
4 | Proposed system should be safe for paediatric patients and have smooth, rounded and sealed surfaces wherever practicable.
5 | Proposed system should have no exposed sharp edges, burrs, pinch points or hazardous projections.
6 | Proposed system should be constructed from durable, non-toxic and cleanable materials.
7 | Proposed system should be resistant to commonly used hospital cleaning/disinfection products.
8 | Applicable items in the proposed system must comply with ASTM F963-23 which addresses hazards associated with toys, including material quality, flammability, accessible edges, projections, stability and overload requirements.
9 | Occupational therapy items should include (but not limited to): • Pegboard Sets • Fine Motor Manipulation Set • Hand Exercise Set • Lacing/Threading Board • Tweezer/Manipulation Activity Set • Construction/Assembly Set
10 | Proposed system should include a storage solution for sensory and occupational therapy equipment: • Heavy-duty shelving. • Rounded/protected edges. • Lockable cabinets for small components where required. • Open storage for frequently used large equipment. • Labelled storage locations. • Adequate ventilation where required. • Easy-clean surfaces. • Storage capacity sized for the complete equipment schedule.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .

## Pneumatic Gym Equipment
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Pneumatic Upper and Lower Body Multifunction Machine | 2 | Comply | Supporting Documents Ref. | by Hospital
2 | Pneumatic Pulley System Free Standing | 2
3 | Pneumatic Leg Press | 1
4 | Pneumatic Leg Extension + Curl | 2
5 | Pneumatic Back and Abdominal Trainer | 2
6 | Pneumatic Total Hip Trainer | 2
7 | Pneumatic Chest Press Machine | 2
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Pneumatic Upper and Lower Body Multifunction Machine | 2
Please specify make, model, and origin of the machine.
1.1 | Unit should provide pneumatic resistance.
1.2 | Please specify dimenstions, weight, power consumption and heat generation.
1.3 | Please specify if unit requires external compressor.
1.4 | Unit should be able to train five exercises: legs are exercised with Leg Extension and Curl, arms and shoulders with Push Up, Pull Down and Chest Press exercises.
1.5 | Unit should have hand grips.
1.6 | Unit should have safety belt strap.
1.7 | Unit should have adjustable back support.
1.8 | Unit leg lever should be adjustable to accommodate for different patient limb lengths.
1.9 | Unit should have leg straps to secure patient thighs during treatment.
1.1 | Unit should have range of motion limiters with start and end positions.
1.11 | Unit should come with a color touch screen to adjust treatment parameters.
1.12 | Unit should allow for small incremental increases of weight of 100g.
1.13 | Unit resistance should start at a load near 0g.
1.14 | FDA approved / CE Marked.
1.15 | Offered System must comply with Electrical safety IEC 6060-1.
2 | Pneumatic Pulley System Free Standing | 2
Please specify make, model, and origin of the machine.
2.1 | Unit should provide pneumatic resistance.
2.2 | Please specify dimensions, weight, power consumption and heat generation.
2.3 | Please specify if unit requires external compressor.
2.4 | Pulley should have adjustable support handles.
2.5 | Unit should come with pulley hand grips.
2.6 | Unit should come with pulley leg straps.
2.7 | Unit should come with a color touch screen to adjust treatment parameters.
2.8 | Unit should allow for small incremental increases of weight of 100g.
2.9 | Unit resistance should start at a load near 0g.
2.10 | FDA approved / CE Marked.
2.11 | Offered System must comply with Electrical safety IEC 6060-1.
3 | Pneumatic Leg Press | 1
Please specify make, model, and origin of the machine.
3.1 | Unit should provide pneumatic resistance.
3.2 | Please specify dimensions, weight, power consumption and heat generation.
3.3 | Please specify if unit requires external compressor.
3.4 | Unit should have hand grips.
3.5 | Unit should have adjustable seat.
3.6 | Unit should have foot straps to secure patient feet during treatment.
3.7 | Unit should have locking mechanism to adjust range of movement and set desired starting point of flexion.
3.8 | Unit should be able to train one leg at a time.
3.9 | Unit should come with a color touch screen to adjust treatment parameters.
3.10 | Unit should allow for small incremental increases of weight of 100g.
3.11 | Unit resistance should start at a load near 0g.
3.12 | FDA approved / CE Marked.
3.13 | Offered System must comply with Electrical safety IEC 6060-1.
4 | Pneumatic Leg Extension + Curl | 2
Please specify make, model, and origin of the machine.
4.1 | Unit should provide pneumatic resistance.
4.2 | Please specify dimensions, weight, power consumption and heat generation.
4.3 | Please specify if unit requires external compressor.
4.4 | Unit should be able to train patient leg extension and leg curl in a single machine (opposing muscle groups).
4.5 | Unit should have hand grips.
4.6 | Unit should have safety belt strap.
4.7 | Unit should have adjustable back stupport.
4.8 | Unit leg lever should be adjustable to accommodate for different patient limb lengths.
4.9 | Unit should have leg straps to secure patient thighs during treatment.
4.10 | Unit should have range limiters with start and end positions.
4.11 | Unit should come with a color touch screen to adjust treatment parameters.
4.12 | Unit should allow for small incremental increases of weight of 100g.
4.13 | Unit resistance should start at a load near 0g.
4.14 | FDA approved / CE Marked.
4.15 | Offered System must comply with Electrical safety IEC 6060-1.
5 | Pneumatic Back and Abdominal Trainer | 2
Please specify make, model, and origin of the machine.
5.1 | Unit should provide pneumatic resistance.
5.2 | Please specify dimensions, weight, power consumption and heat generation.
5.3 | Please specify if unit requires external compressor.
5.4 | Unit should be able to train patient back and abdominal muscles in a single machine (opposing muscle groups).
5.5 | Unit should come with a color touch screen to adjust treatment parameters.
5.6 | Unit should have adjustable seat height.
5.7 | Unit should have adjustable lever arm.
5.8 | Unit should allow for small incremental increases of weight of 100g.
5.9 | Unit resistance should start at a load near 0g.
5.10 | FDA approved / CE Marked.
5.11 | Offered System must comply with Electrical safety IEC 6060-1.
6 | Pneumatic Total Hip Trainer | 2
Please specify make, model, and origin of the machine.
6.1 | Unit should provide pneumatic resistance.
6.2 | Please specify dimensions, weight, power consumption and heat generation.
6.3 | Please specify if unit requires external compressor.
6.4 | Unit should be able to train patient adductor and abductor muscles for hips and legs in a single machine (opposing muscle groups).
6.5 | Unit should have hand grips.
6.6 | Unit should have range limiters with start and end positions.
6.7 | Unit should come with a color touch screen to adjust treatment parameters.
6.8 | Unit should allow for small incremental increases of weight of 100g.
6.9 | Unit resistance should start at a load near 0g.
6.10 | FDA approved / CE Marked.
6.11 | Offered System must comply with Electrical safety IEC 6060-1.
7 | Pneumatic Chest Press Machine | 2
Please specify make, model, and origin of the machine.
7.1 | Unit should provide pneumatic resistance.
7.2 | Please specify dimensions, weight, power consumption and heat generation.
7.3 | Please specify if unit requires external compressor.
7.4 | Unit should be able to train patient chest and arms in a chest press exercise motion.
7.5 | Unit should come with a color touch screen to adjust treatment parameters.
7.6 | Unit should have multigrip handles.
7.7 | Unit should have range limiters with start and end positions.
7.8 | Unit should have adjustable seat height.
7.9 | Unit should allow for small incremental increases of weight of 100g.
7.10 | Unit resistance should start at a load near 0g.
7.11 | FDA approved / CE Marked.
7.12 | Offered System must comply with Electrical safety IEC 6060-1.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .

## Physio Machines
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Class 4 Laser Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
2 | Pelvic Stimulator Machine | 1
3 | Cold pack Unit | 1
4 | Hot pack Unit | 1
5 | Combination Therapy Machine 4-Channel | 3
6 | Combination Therapy Machine 2-Channel | 2
7 | Patient Hoist Lifter | 2
8 | Bobath Bed | 1
9 | 3 Section Treatment Bed | 10
10 | Bariatric Wheelchair | 2
11 | Adult Wheelchair | 6
12 | Upper and Lower Limb Exerciser | 1
13 | Treadmill | 2
14 | Compressive Cryotherapy Unit | 1
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Laser Machine Class 4 | 1
Please specify make, model, and origin of the machine.
1.1 | Unit should be class 4 laser.
1.2 | Unit should have control buttons on the handpiece.
1.3 | Unit handpiece should have stationary arm to allow for hands-free treatment.
1.4 | Unit should come with all laser safety accessories including 3 safety goggles per unit.
1.5 | Pulse range from 1-20,000 Hz.
1.6 | Unit should have at least 4 wavelengths (660nm, 800nm, 905nm, 970nm) which can be emitted simultaneously or in any combination.
1.7 | Unit should have pre-programmed protocols for treatments.
1.8 | Protocol selection by size, body part, skin tone, pain and chronicity level automatically adjust output treatment parameters for increased safety and effectiveness.
1.9 | Variable zoom from 1 to 5 cm2
1.1 | Intense Super Pulse (ISP) 20 W
1.11 | Continuous wave (CW) 15 W
1.12 | Unit should have battery backup of at least 1 hour.
1.13 | USB port for quick export of patient history to computer
1.14 | Unit should have large color touchscreen to adjust treatment parameters.
1.15 | Unit should be mobile and come with a trolley with wheels.
1.16 | FDA approved / CE Marked .
1.17 | Offered System must comply with Electrical safety IEC 6060-1.
1.18 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Pelvic Stimulator Machine | 1
Please specify make, model, and origin of the machine.
2.1 | Non-invasive neuromuscular electrical stimulation for pelvic-floor rehabilitation under trained clinician supervision
2.2 | Frequency range: Adjustable, with a minimum range covering approximately 5–100 Hz
2.3 | Pulse width: Adjustable, with a minimum range covering approximately 50–400 μs
2.4 | Treatment protocols: User-programmable treatment protocols and/or predefined pelvic-floor rehabilitation programs
2.5 | Unit should have digital display showing treatment parameters, intensity, treatment time and relevant biofeedback information.
2.6 | Adjustable treatment timer with automatic termination at completion.
2.7 | FDA approved / CE Marked.
2.8 | Offered System must comply with Electrical safety IEC 6060-1.
2.9 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
3 | Cold pack Unit | 1
Please specify make, model, and origin of the machine.
3.1 | Please specify unit dimensions.
3.2 | Closed-cell foam insulation for efficiency.
3.3 | Unit made of stainless steel.
3.4 | Unit should have heavy-duty compressor.
3.5 | Unit should be mobile with swivel castor wheels for silent, friction-free movement.
3.6 | Unit should have drain valve for easy cleaning and defrosting.
3.7 | Unit capacity should be able to chill at least 12 cold packs at a time.
3.8 | Unit should come with included cold packs.
3.9 | FDA approved / CE Marked.
3.10 | Offered System must comply with Electrical safety IEC 6060-1.
4 | Hot pack Unit | 1
Please specify make, model, and origin of the machine.
4.1 | Please specify unit dimensions.
4.2 | Unit made of stainless steel.
4.3 | Unit should be mobile with swivel castor wheels.
4.4 | Unit should have a capacity of at least 135 liters.
4.5 | Unit should be thermostatically controlled to ensure the ideal therapeutic temperature for hot packs.
4.6 | Unit should have an overtemperature switch that will switch off the unit in case of overheating.
4.7 | Unit should have drain valve for easy cleaning and draining.
4.8 | Unit should come with included hot packs
4.9 | FDA approved / CE Marked.
4.10 | Offered System must comply with Electrical safety IEC 6060-1.
5 | Combination Therapy Machine 4-channel | 3
Please specify make, model, and origin of the machine.
5.1 | Unit should accommodate 4 channels.
5.2 | Unit should have the following treatment options: • Electro therapy, 2 and 4 poles, 2 channels completely independent. • S-D curve diagnostic programs. • Ultrasound therapy. • Simultaneous therapy (2 different indications treated simultaneously by using electro and ultrasound therapy). • Combination therapy (treating one injury simultaneously using a combination of electro and ultrasound therapy).
5.3 | Unit should come with 1 large and 1 small ultrasound applicator.
5.4 | Unit should come with a pen electrode for electrotherapy treatment.
5.5 | Ultrasound applicators should have multifrequency head (1 and 3 MHz).
5.6 | Ultrasound should have continuous and pulsed mode (10–20–30–40–50-100%).
5.7 | Unit should have the following unidirectional currents: • Direct current • Rectangular pulse • 2-5 current (Ultra Reiz) • Triangular pulse • MF rectangular pulse • Iontophoresis- MF rectangular pulse • Iontophoresis- direct current
5.8 | Unit should have the following Diadynamic currents: • MF • DF • RS • CP • LP
5.9 | Unit should have the following TENS currents: • Conventional TENS • Low frequency TENS • Burst TENS • High frequency TENS • Random Frequency TENS • Han Stim (via painrelief)
5.10 | Unit should have the following NMES currents: • Rectangular surge • Triangular surge • Biphasic surge • Intrapuls interval surge • Russian stimulation • 2-pole MF surge • Isoplanar vector field surge
5.11 | Unit should have the following Interferential currents: • 2-pole Medium Frequency • Isoplanar vector field • Dipole vector field • Classical interferential
5.12 | Unit should have micro current, micro current modulated, and microcurrent surge.
5.13 | Unit should have high voltage (HVPC) and high voltage surge currents.
5.14 | Unit should be mobile and come with a trolley with wheels.
5.15 | Unit should have a color touch screen to adjust treatment parameters.
5.16 | FDA approved / CE Marked.
5.17 | Offered System must comply with Electrical safety IEC 6060-1.
5.18 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
6 | Combination Therapy Machine 2-channel | 2
Please specify make, model, and origin of the machine.
6.1 | Unit should accommodate 2 channels.
6.2 | Unit should have the following treatment options: • Electro therapy, 2 and 4 poles, 2 channels completely independent. • S-D curve diagnostic programs. • Ultrasound therapy. • Simultaneous therapy (2 different indications treated simultaneously by using electro and ultrasound therapy). • Combination therapy (treating one injury simultaneously using a combination of electro and ultrasound therapy).
6.3 | Unit should come with 1 large and 1 small ultrasound applicator.
6.4 | Unit should come with a pen electrode for electrotherapy treatment.
6.5 | Ultrasound applicators should have multifrequency head (1 and 3 MHz).
6.6 | Ultrasound should have continuous and pulsed mode (10–20–30–40–50-100%).
6.7 | Unit should have the following unidirectional currents: • Direct current • Rectangular pulse • 2-5 current (Ultra Reiz) • Triangular pulse • MF rectangular pulse • Iontophoresis- MF rectangular pulse • Iontophoresis- direct current
6.8 | Unit should have the following Diadynamic currents: • MF • DF • RS • CP • LP
6.9 | Unit should have the following TENS currents: • Conventional TENS • Low frequency TENS • Burst TENS • High frequency TENS • Random Frequency TENS • Han Stim (via painrelief)
6.10 | Unit should have the following NMES currents: • Rectangular surge • Triangular surge • Biphasic surge • Intrapuls interval surge • Russian stimulation • 2-pole MF surge • Isoplanar vector field surge
6.11 | Unit should have the following Interferential currents: • 2-pole Medium Frequency • Isoplanar vector field • Dipole vector field • Classical interferential
6.12 | Unit should have micro current, micro current modulated, and microcurrent surge.
6.13 | Unit should have high voltage (HVPC) and high voltage surge currents.
6.14 | Unit should be mobile and come with a trolley with wheels.
6.15 | Unit should have a color touch screen to adjust treatment parameters.
6.16 | FDA approved / CE Marked.
6.17 | Offered System must comply with Electrical safety IEC 6060-1.
6.18 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
7 | Patient Hoist Lifter | 2
Please specify make, model, and origin of the machine.
7.1 | Unit should have a remote control to control the lifter.
7.2 | Unit should have a safe working load of 300kg.
7.3 | Unit maximum lifting height should be at least 191 cm.
7.4 | Unit should have a weighing scale integrated in the lifter.
7.5 | Unit should have an emergency stop button.
7.6 | Unit should have a rechargeable battery.
7.7 | Unit should come with resuable patient sling.
7.8 | FDA approved / CE Marked.
7.9 | Offered System must comply with Electrical safety IEC 6060-1.
7.10 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
8 | Bobath Bed | 1
Please specify make, model, and origin of the machine.
8.1 | Unit should have 2 sections.
8.2 | Unit dimensions should be 208 cm length and 120 cm width.
8.3 | Head section angle should be adjustable.
8.4 | Unit should be electric and have a footswitch to adjust the height.
8.5 | Unit should have retractable castors at both ends to allow for easy movement and anchoring.
8.6 | FDA approved / CE Marked.
8.7 | Offered System must comply with Electrical safety IEC 6060-1.
9 | 3 Section Treatment Bed | 10
Please specify make, model, and origin of the machine.
9.1 | Unit should have 3 sections.
9.2 | Unit dimensions should be 195 cm length and 70 cm width.
9.3 | Unit should be electric and have a footswitch to adjust the height.
9.4 | Unit head section should have a hole for the patient's face.
9.5 | Unit should have retractable castors at both ends to allow for easy movement and anchoring.
9.6 | FDA approved / CE Marked.
9.7 | Offered System must comply with Electrical safety IEC 6060-1.
10 | Bariatric Wheelchair | 2
Please specify make, model, and origin of the machine.
10.1 | Unit should be able to accommodate bariatric patients weighing up to 300 kg.
10.2 | Seat width should be at least 28 inches.
10.3 | Unit should be made of heavy duty material.
10.4 | Unit should have armrests.
10.5 | Unit should have adjustable leg rests.
10.6 | Unit should have wheel brakes.
10.7 | FDA approved / CE Marked.
11 | Adult Wheelchair | 6
Please specify make, model, and origin of the machine.
11.1 | Unit should be able to accommodate patients weighing up to 115 kg.
11.2 | Seat width should be at least 16 inches.
11.3 | Unit backrest should be able to fully recline to a flat position.
11.4 | Unit should have armrests.
11.5 | Unit should have adjustable leg rests.
11.6 | Unit should have wheel brakes.
11.7 | FDA approved / CE Marked.
12 | Upper and Lower Limb Exerciser | 1
Please specify make, model, and origin of the machine.
12.1 | Unit should be able to train upper and lower limbs
12.2 | Unit should feature 7-inch color touch screen.
12.3 | Unit should feature soft grip foot shells with fastening straps.
12.4 | Unit should feature soft grip handles.
12.5 | Height adjustment of leg or arm/upper body trainer can be adjusted without tools.
12.6 | Pedal axle of the leg trainer is height adjustable from 30 cm to 45 cm.
12.7 | Pedal axle of the arm trainer is height adjustable from 90 cm to 105 cm.
12.8 | Unit should feature pedal radius adjustment.
12.9 | Closed plastic body (PC/ABS), easy to clean and disinfect.
12.10 | Easily movable with large transport rollers
12.11 | FDA approved / CE Marked.
12.12 | Offered System must comply with Electrical safety IEC 6060-1.
13 | Treadmill | 2
Please specify make, model, and origin of the machine.
13.1 | Unit should have adjustable speed from at least 1 km/h to 18 km/h.
13.2 | Unit should have adjustable incline angle (15 levels of incline).
13.3 | Unit should have emergency stop button.
13.4 | Unit should have preset programs.
13.5 | Unit should have LCD display to display speed, time, distance, incline level.
13.6 | FDA approved / CE Marked.
13.7 | Offered System must comply with Electrical safety IEC 6060-1.
14 | Compressive Cryotherapy Unit | 1
Please specify make, model, and origin of the machine.
14.1 | Unit should have multiple treatment programs.
14.2 | Unit should come with all standard wraps for treatment of body regions including knee, ankle, articulated knee, leg, shoulder, hand/wrist, elbow, back, hip.
14.3 | Unit should be easy to clean and have a rinse program.
14.4 | FDA approved / CE Marked.
14.5 | Offered System must comply with Electrical safety IEC 6060-1.
Accessories
Supplied with all standard accessories.
Consumables
Supplied with all standard consumables.
Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warrenty .